## Slide 1: Dr. Eric Chatfield Chatfield Technical Consulting Limited 2071 Dickson Road Mississauga, Ontario, Canada L5B 1Y8
Chatfield is Lead Author of:
  Analytical Method for Determination of Asbestos Fibers in Water, EPA-600/4-83-043, now identified as EPA Method 100.1.
  ISO 10312 Ambient air — Determination of asbestos fibres — Direct transfer transmission electron microscopy method.
  ISO 13794 Ambient air — Determination of asbestos fibres — Indirect-transfer transmission electron microscopy method.
  ISO 22262-1 Air quality — Bulk materials -- Part 1: Sampling and qualitative determination of asbestos in commercial bulk materials.
  ISO 22262-2 Air quality -- Bulk materials -- Part 2: Quantitative determination of asbestos by gravimetric and microscopical methods.
Dr. Chatfield has provided testimony for both defendants and plaintiffs in asbestos litigation.
He is also a member of the Science Advisory Board of the National Stone, Sand and Gravel Association.

## Slide 2: Comments on the Executive Summary, Dated January 6, 2020
The statement “Published methods for analysis of bulk materials were not intended to determine the presence of asbestos in products at less than 1% concentration” is incorrect. ISO22262-1:2012 and ISO22262-2:2014 were specifically written to accommodate national standards down to the U.K. regulation of “any asbestos”.
On Page 4, Footnote 11, references in support of including EMP’s between 0.5 µm and 5 µm in the measurements are somewhat selective:
  Publications that report zero potency or minimal potency for EMPs in this dimensional range have been overlooked e.g. Berman et al. (1995), Berman and Crump (2008), Roggli (2015).
  The 2002 and 2005 Suzuki et al. papers do not conform to modern analytical TEM methods. The fibers reported were measured on photographic prints, and these fibers appear to be classified only on the basis of morphology. Moreover, fibers <0.5 µm were included in the counts; 0.5 µm has been accepted since 1995 in ISO TEM methods as the minimum length for reliable identification and detection.
  ANSES (2015) is erroneously cited in support of including EMP’s between 0.5 µm and 5µm in the measurements. In 12 – Recommendations, ANSES (2015) actually states the opposite (translated from the full document that is in French only):
    That only EMPs having a length greater than 5 µm and a diameter of less than 3 µm (L > 5 µm, D < 3 µm) be considered in natural materials.
    That air and material analysis reports mention the presence of EMPs shorter than 5 µm and having a diameter of less than 3 µm (L ≤ 5 µm, D < 3 µm) to permit traceability.
There is no basis for quantitative risk assessment for EMPs between 0.5 µm and 5 µm.

## Slide 3: Recommendations
A draft revision of ISO 22262-2 has been prepared to include reporting of numerical EMP concentrations in talc and other mineral powders. The revised ISO 22262-2 defers to ISO 13794 for liquid filtration methods and all other procedures for TEM analysis of particles collected on a filter.
A TEM count of EMPs for all lengths ≥0.5 µm usually means that EMPs >5 µm, although present, may not be detected. If short EMPs are to be counted, a dual-count approach is needed: one count at 20,000x for EMPs ≥0.5 µm and another at 10,000x for EMPs >5 µm.
For thinner EMPs, a higher magnification of ~60,000 is needed for more accurate measurement of fiber widths.
Reliable determination of mass percent is possible, but it requires a specific counting protocol. Mass is the only invariant measurement. For numerical EMP concentrations per gram, a target analytical sensitivity must be defined. What is the level of concern in terms of EMPs/gram or mass percent?
With respect to the minimum length of EMP that should be counted, examine the process by which EPA defined the MCL for drinking water in terms of fibers longer than 10 µm.
Possible Areas of Research on Concentration Methods
  For separation of chrysotile from talc, methods based on the difference between the electrical charge on talc and chrysotile may be effective. A starting point would be Melton et al. (EPA 600/4-78-066, Development of a Rapid Analytical Method for Determining Asbestos in Water).
  For separation of amphibole from talc, density separation works for amphiboles other than low-iron anthophyllite. There does not currently appear to be any other approach for separation of amphibole.
In determining the sizes of EMP to include in measurements and potential regulation, take note of the results of animal studies.

## Slide 4
JAWE 431 Anthophyllite
0% Mesotheliomas
Davis et al. (1991)
Aierken et al. (2014)
Korean Tremolite
97% Mesotheliomas
Shinness Tremolite
5.6% Mesotheliomas
JAWE 531 Tremolite
96% Mesotheliomas

## Slide 5
Combined Measurements
from Analyses of 30 Talc
Products, Performed by a
Laboratory for Litigation
in 2017.
Tremolite, Richterite and
Ferro-anthophyllite
Aierken et al. (2014) JAWE 431/411 Anthophyllite
“in this preparation was not carcinogenic to mesothelial cells…”
Davis et al. (1991) Shinness Tremolite
“…almost certainly …harmless to human beings”