Oncology Pilot Clinical Trial Assay Template:

Immunohistochemistry Tests

For Collection of Performance Validation Data

SCOPE: This template is intended for use by oncology drug sponsors in the Food and Drug Administration’s (FDA) voluntary pilot program described in FDA’s guidance document “Oncology Drug Products Used with Certain In Vitro Diagnostic Tests: Pilot Program.” In this and other templates, FDA is providing recommendations regarding data and information that should be submitted to FDA regarding clinical trial assays (CTAs) used in the pivotal clinical trial(s) for oncology drug product(s) for the pilot program. This template is specific for CTAs that use Immunohistochemistry (IHC) technology. Templates for CTAs that use other technologies can be found on FDA’s website.

The CTA templates are intended to help oncology drug sponsors and CTA developers collect and provide validation information and performance characteristics for the CTAs, but alternative approaches can be used. To discuss an alternative approach, please contact OncologyPilotCDRH@fda.hhs.gov.

This template reflects FDA’s current thinking on the topic, and should be viewed only as recommendations, unless specific regulatory or statutory requirements are cited. The use of the word should, means that something is suggested or recommended, but not required.

As described in FDA’s guidance document, oncology drug sponsors interested in participating in the pilot program should submit a statement of interest to their Investigational New Drug (IND) application, New Drug Applications (NDA), or Biologic License Applications (BLA), as appropriate. The statement of interest should include a statement affirming the oncology drug sponsor’s commitment to provide the information recommended in this template for CTAs used in their pivotal clinical trial(s) and the additional information described in the guidance. Upon receipt of the statement of interest, FDA will follow up with no more than 9 sponsors to request specific information, including the information identified below, to enable FDA to make a decision concerning acceptance into the pilot, based on evaluation of the factors outlined in FDA’s guidance document and provide written feedback that either accepts or rejects the drug product for the pilot program.  This template should be used to provide the information recommended in this template only when requested by FDA.

GENERAL INFORMATION ABOUT THIS TEMPLATE

- When requested by FDA, oncology drug sponsors who have submitted the statement of interest described above should complete the information below, as applicable, using the blue fillable fields. Hot links are provided to navigate to and from sections where data entry into tables is recommended; these data tables are provided in the Appendix.

- Alternatively, oncology drug sponsors have the option of providing the recommended information by submitting documentation from the laboratory that performed the CTA such as the standard operating procedure (SOP) and a summary of test validation.

General Laboratory Information

- Laboratory name and address: Enter laboratory’s name and address

- Laboratory contact name and email: Enter person’s name and email address

- Test name: Enter test name

General Test Information

Test information

Please provide a general description of the test including:

  - Is the test/CTA commercially available? ☐ Yes; ☐ No

    - if yes, provide the following information:

      - name of the kit and manufacturer: Enter kit name(s) and manufacturer(s) here

      - if any modifications to the kit were made, describe them: Summarize here

    - if not, provide the following information:

      - analyte(s) (e.g., PD-L1 protein): Enter analyte(s)

      - test method, including:

        - antibody clone (e.g., rabbit monoclonal anti-PD-L1 clone SP142): Enter antibody clone

        - format (e.g., single antibody, multiplex): Enter format

        - type of measurement: Choose an item

        - specimen type (e.g., formalin-fixed paraffin-embedded [FFPE] breast tumor tissue): Enter specimen type

  - Detection system: Enter detection system

  - Staining instrument or platform: Enter staining instrument

  - Control source material (e.g., FFPE tissue and /or cell lines) and control description: Enter source and description

Scoring methodology & clinical cut-off

- Provide the instructions for slide review and interpretation of results: Enter instructions

- Please identify the components of the scoring algorithm used, including:

- staining intensity: Enter staining intensity

- and/or proportion of cells staining: Enter proportion staining

- staining localization (e.g., nuclear, cytoplasmic, membranous): Enter locations

- staining in tumor/non-tumor elements and areas of the specimen: Enter elements and areas

- Briefly summarize the prespecified clinical cut-off that was used to enroll subjects in the clinical trial: Summarize here

Validation Studies

- Study bias

Please provide the following information regarding potential study bias:

- Separate sets of specimens/slides were used for training and validation?

☐ Yes; ☐ No

- Study slides were read in an unbiased manner (i.e., readers were blinded to the specific slide ID) and to the specific analytical validation parameter (e.g., between-day, within-run, use of an adequate wash-out period between sets of slides, etc.):

☐ Yes; ☐ No

Primary IHC antibody characterization

  - If the characterization of the IHC antibody was established by the antibody manufacturer, please provide the information from the manufacturer (e.g., method used such as data from Western blot analysis) as an attachment. If it was established by another third party and published in the scientific literature, please provide the publications where the IHC antibody clone used in the test was characterized: PubMed references

  - If performed in the laboratory, please provide the method performed to demonstrate antibody specificity to antigen (e.g., Western blot analysis): Enter method

Analytical accuracy – concordance

Analytical accuracy of the CTA is determined relative to a reference method or validated comparator method (orthogonal method). If a comparator method is used, it should be measuring the same analyte with similar sensitivity and be using the same intended use clinical specimens as the candidate test (i.e., the CTA). Well characterized samples should be tested with both the CTA and the comparator method.

Please provide the following information:

- Briefly summarize the study design, including the statistical/data analysis methods used to determine analytical accuracy: Summarize here

- The name of the comparator method: Method name

- The technology of the comparator method: Technology

- If a laboratory developed test was used as a comparator:

    - please provide the laboratory name and address: Enter laboratory’s name and address

    - please provide the stated accuracy of the comparator method: Enter accuracy

- Analytical accuracy study performed in-house? ☐ Yes; ☐ No

- Describe the range of clinical test results (i.e., expression levels) for the accuracy study samples: Describe range

- Positive percent agreement (PPA) and negative percent agreement (NPA) between the CTA and the comparator method should be calculated. Provide concordance summary data in Table 1.

Analytical specificity

Please provide information on immunoreactivity in human tissues (through a tour of tumor and tour of body studies).

Please provide information demonstrating that:

- The immune reactivity of the IHC test (i.e., primary antibody with an acceptable secondary antibody, detection, and other ancillary reagents) was validated by testing tissues or intact cells that do and do not contain the target analyte in the same specimen matrix found in the clinical trial: Enter immune reactivity information

- For normal tissue and tumor tissue the staining was as expected: Enter data here

Precision

Precision studies are performed to evaluate sources of variation in the test procedure to determine the repeatability and reproducibility of the test.

Please provide the following information:

  - A brief study design (e.g., full factorial): Summarize here

- Runs: Enter number

- Days: Enter number

- Instruments: Enter number

- Reagent lots: Enter number

- Operators: Enter number

- Pathologist readers: Enter number

- Number of replicates per sample: Enter number

- The number of antibody lots used: Enter # of lots here

- Please provide a summary in Table 2 of the results with positive percent agreement (PPA) and negative percent agreement (NPA).

Stability

Stability studies are conducted to support storage conditions, including the duration of storage, for specimens and stored intermediate products, as applicable.

Please state the stability for:

- Primary specimens: Enter condition(s) and duration(s)

- Intermediate specimen products (e.g., unstained tissue sections mounted on a glass slide): Enter storage condition(s) and duration(s)

Appendix: Data Tables

Table 1: Summary of the analytical concordance between CTA IHC (CTA) and comparator method

|  |  | Comparator method |  |
| --- | --- | --- | --- |
| CTA | Positive | Negative | Total |
| Positive | Enter data | Enter data | Enter data |
| Negative | Enter data | Enter data | Enter data |
| Total | Enter data | Enter data | Enter data |
| Agreement: | PPA: Enter data NPA: Enter data |  |  |

To return to “Analytical accuracy – concordance” click here

Table 2: Summary of precision

| Precision study parameter | Range around the cut-off | Number positives / number expected | Number negatives / number expected | % Agreement |
| --- | --- | --- | --- | --- |
| Within-day | Enter data | Enter data | Enter data | PPA: Enter data NPA: Enter data |
| Between-day | Enter data | Enter data | Enter data | PPA: Enter data NPA: Enter data |
| Between-reader (visual) | Enter data | Enter data | Enter data | PPA: Enter data NPA: Enter data |
| Between instrument (the stainer) | Enter data | Enter data | Enter data | PPA: Enter data NPA: Enter data |

To return to “Precision” click here