## GUDID_Data_Elements
Data Element 6 | Description | Data Entry Notes | Edit Rules After Grace Period 1 | Required in Database? 2 | Data Type & Length 3 | Entry List of Values (LOV) | New DI Trigger | Publicly Released?
According to 830.330(b), The labeler of a device shall submit to FDA an update to the information required by § 830.310 whenever the information changes. The updated information must be submitted no later than the date a device is first labeled with the changed information. If the information does not appear on the label of a device, the updated information must be submitted within 10 business days of the change.
When a GUDID attribute appears in the medical device labeling, the values submitted to the GUDID should match the value in the labeling.
Device Information
Device Identifier (DI) Information
Issuing Agency | Organization accredited by FDA to operate a system for the issuance of UDIs. | Choose a value from the drop down List of Values (LOV). | None | Required | NA | GS1; HIBCC; ICCBBA | Yes | Yes
Primary DI Number | The Primary Device Identifier (DI) is the DI portion of the UDI placed on the lowest package level of a device that is required to meet UDI label requirements. If the device is not packaged, the UDI may be on the device itself, thereby satisfying both the UDI label and the direct mark (DM) requirement if the UDI is intended to be permanent. The primary DI is the main (primary) lookup for a medical device and meets the requirements to uniquely identify a device through its distribution and use. | Enter the Device Identifier (DI) Number. If using GS1 as an issuing agency, ensure the check digit is correct and valid as per the GS1 guidelines. If using HIBCC or ICCBBA as an issuing agency, do not include the check digit/character as part of your DI. Data type and field length are determined by the individual Issuing Agency structure. IF Device IS Packaged Primary DI = DI portion of the UDI on the lowest package level OR IF Device is NOT Packaged Primary DI = DI on the device itself, could be DM DI GS1: Numeric (Num.), DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanumeric (Alphanum.), with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters | None | Required | Type: Num. or Alphanum. Length: min-6, max-23* *defined by Issuing Agency structure. | NA | Yes | Yes
Device Count | Number of medical devices in the base package. | Enter the number of devices. Example: Base Package = Box of 100 gloves; Primary DI = 101; Device Count = 100. | None | Required | Type: Num. Length: 7 | NA | Yes | Yes
Unit of Use DI Number | An unmarked identifier assigned to an individual medical device when a UDI is not labeled on the individual device at the level of its unit of use. Its purpose is to associate the use of a device to/on a patient. The Unit of Use DI does not appear on the label of the device. | Enter the Unit of Use DI Number. Must be from same Issuing Agency as Primary DI. Unit of Use DI is an identifier used by hospital staff and Materials Management to account for a single device when the UDI is labeled on a higher level of packaging. The Unit of Use DI does not appear on the label. Data type and field length are determined by the individual Issuing Agency structure. GS1: Num., DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanum., with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters If Device Count = 1, cannot add Unit of Use DI Number. | Edit | Conditionally Required* *If Device Count >1. | Type: Num. or Alphanum. Length: min-6, max-23* *defined by Issuing Agency structure. | NA | No | Yes
Labeler DUNS Number | Business number Issued by Dun & Bradstreet (D&B) that is used to associate the Labeler (Company) name and address to a given version of model of a device in GUDID. | Choose appropriate DUNS Number from drop down LOV. To ensure data consistency for the GUDID, DUNS number submitted to the GUDID should associate to the company name that appears on the device label; ideally the address associated with the DUNS number should also match the address on the device label, but since address is not displayed to the GUDID public user, this is not a requirement for data consistency. All edits to information connected to the Labeler DUNS Number must be done through Dun & Bradstreet. No edits of DUNS information will be permitted in the GUDID. | Edit* *Other Labeler DUNS listed to your GUDID account can be selected. No edits of DUNS info will be permitted. | Required | NA | Labeler DUNS LOV | No | Yes
Company Name | Company name associated with the labeler DUNS Number entered in the DI Record. | Auto populated based on the Labeler DUNS Number The labeler company name submitted to the GUDID should match the company name on the device label. | NA | Auto Populated | NA | NA | No | Yes
Company Physical Address | Company physical address associated with the labeler DUNS Number entered in the DI Record. | Auto populated based on the Labeler DUNS Number Ideally, this address should match the labeler address as shown on the device label but since this data element is not displayed to the GUDID public user, this is not a requirement for data consistency. | NA | Auto Populated | NA | NA | No | No
Brand Name | A brand name is a trade/proprietary name assigned by the device labeler, and under which the device is sold, distinguished from other similar devices, and recognized by the user or purchaser. A brand name is often registered and/or protected by a trademark. | Enter the brand name specific to this version/model of the device. You may include the ® and/or TM symbol, if applicable. If there is no brand name for your device then enter NA or the name commonly used to refer to the device or family of devices. Information entered in this field should be greater than one character in length, and a name that refers to the device must be entered. | None | Required | Type: Alphanum. Length: 80 | NA | Yes | Yes
Version or Model | The version or model found on the device label or accompanying packaging used to identify a category or design of a device. The version or model identifies all devices that have specifications, performance, size, and composition within limits set by the labeler. | Enter the Version or Model. Version/Model can be any distinguishing string of letters and/or numbers. If the device does not have a version or model, enter a concept that can be used to identify all devices that have specifications, performance, size, and composition within limits set by the labeler. | None | Required | Type: Alphanum. Length: 80 | NA | Yes | Yes
Catalog Number | The catalog, reference, reorder or product number used by the labeler for business and clinical transactions to identify a particular product; this number may be found on the device label or labeling. | The catalog number is critical for matching UDI to legacy device identification data used in the item masters, inventory systems, EHRs, and registries. The DI of the UDI will replace the catalog number for recalls and adverse event reports in the future. To populate this field: Use the catalog, reference, reorder or product number on the device label. If there is no catalog, reference, reorder or product number use the number reported in a business/clinical transaction or that would be reported in case of a recall or adverse event. | Add Delete Edit | Recommended - Critical for identification and adoption | Type: Alphanum. Length: 80 | NA | No | Yes
Device Description | Additional information about the device that is not already captured as a distinct GUDID data element. | Device description should include any description found on the device label, package insert, or approved or cleared indications for use to support user identification. If you have multiple device records with similar information, this field should help users differentiate among the device records. You may include URL links to: the label, instructions for use, device pictures/schematics. The inclusion of a URL is to supplement the description of the device. URLs can change over time. If a URL is included, please keep the URL up to date. Duplicating information already provided as distinct data elements in GUDID, e.g. Brand name, company name, etc. is not necessary. •Note: This field appears in the results summary from searches in AccessGUDID. | Add Delete Edit | Recommended - Critical for identification and adoption | Type: Alphanum. Length: 2000 | NA | No | Yes
Commercial Distribution
DI Record Publish Date | Indicates the date the DI Record is published and available via Public Search. | Choose date from calendar or manually enter date in format yyyy-mm-dd. This date determines the Grace Period, presently set to 30 calendar days. Grace Period starts the day after the DI Record Publish Date. We recommend you set this date to be a future date to allow time to ensure accurate data entry. | None | Required | Type: Date Format (yyyy-mm-dd) Length: 10 | NA | No | Yes
Commercial Distribution End Date | Indicates the date the device is no longer offered for commercial distribution by the labeler on record. See 21 CFR 807.3(b) for exceptions. The device may or may not still be available for purchase in the marketplace. | Choose date from calendar or manually enter date in format yyyy-mm-dd. Do not populate this field for deactivating/deleting records. Please contact the FDA UDI Help Desk at FDA UDI Help Desk. If you do not know an exact commercial distribution end date, please do not populate an estimated date. | Add Delete Edit | Optional | Type: Date Format (yyyy-mm-dd) Length: 10 | NA | No | Yes
Commercial Distribution Status | Indicates whether the device is in commercial distribution as defined under 21 CFR 807.3(b). | Auto populated based on Commercial Distribution End Date. If no Commercial Distribution End Date is entered, the status is 'In Commercial Distribution' | NA | Auto Populated | NA | In Commercial Distribution; Not in Commercial Distribution | No | Yes
Alternative or Additional Identifiers
Direct Marking (DM) |  | Direct Marking (DM) data elements only apply to devices subject to 21 CFR 801.45.
Device Subject to Direct Marking (DM), but Exempt | The device is exempt from Direct Marking requirements under 21 CFR 801.45. | Select checkbox if appropriate. Labeler should select the checkbox “Device Subject to Direct Marking (DM), but Exempt” only if the device: (1) is intended to be used more than once and (2) is intended to be reprocessed before each use, but also (3) meets any one of the exception criteria outlined under 21 CFR 801.45(d). If the device is not required to be directly marked under 21 CFR 801.45(a), then this box should not be checked. | Add Delete Edit | Conditionally Required* *If device is subject to 801.45 | Type: Boolean | NA | No | Yes
DM DI Different from Primary DI | Indicates that the DM DI Number is different than the Primary DI Number. | IF Device IS Packaged AND DM DI ≠ Primary DI on device label THEN Select checkbox DM DI Different from Primary DI and Enter DM DI in DM DI field | Add Delete Edit | Conditionally Required* *If device is subject to 801.45 | Type: Boolean | NA | No | Yes
DM DI Number | An identifier that is permanently marked directly on the medical device. Certain devices are required to be directly marked. Under 21 CFR 801.45(b), the direct mark DI may be identical to or different from the Primary DI. | If the direct mark is identical to the UDI that appears on the device label, do not enter a direct mark DI. If the direct mark is different from the UDI that appears on the device label, the direct mark DI is the DI portion of the direct mark. If a device is not packaged and the UDI is on the device itself and intended to be permanent, direct mark requirements are satisfied, enter the direct mark DI as the primary DI and do not enter a DM DI. The above text is also applicable to devices that are not subject to 21 CFR 801.45, but are voluntarily direct marked with a UDI by the labeler. IF Device IS Packaged AND DM DI = Primary DI on device label THEN Do not enter DM DI ELSE IF DM DI ≠ Primary DI on device label THEN Enter DM DI IF Device IS NOT Packaged AND DI permanently marked on the device (i.e., DM DI) AND DM DI = Primary DI THEN Do not enter DM DI; DM DI is the Primary DI GS1: Num., DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanum., with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters | Add Delete Edit | Conditionally Required* *If device subject to 801.45 and 'DM DI Different from Primary DI' is checked | Type: Num. or Alphanum. Length: min-6, max-23* *defined by Issuing Agency structure | NA | No | Yes
Secondary DI
Issuing Agency | Name of Secondary DI Issuing agency. | Choose a value from the drop down LOV. | None | Optional | NA | GS1; HIBCC; ICCBBA; NDC/NHRIC | No | Yes
Secondary DI Number | An identifier that is an alternate (secondary) lookup for a medical device that is issued from a different issuing agency than the primary DI. Under 21 CFR 830.40(a), only one device identifier from any particular system for the issuance of UDIs may be used to identify a particular version or model of a device. | Enter Secondary DI Number. If your product is labeled with a UDI from more than one issuing agency (for regulatory or marketing reasons), you must choose one issuing agency as the Primary DI and enter the other issuing agency information here, as a Secondary DI. Data type and field length are determined by the individual Issuing Agency structure. GS1: Num., DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanum., with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters NDC: Numeric (Num.), DI must be 10 or 11 digits in 3 segments with 2 hyphens.Valid formats for 10-digit values are XXXX-XXXX-XX, or XXXXX-XXX-XX, or XXXXX-XXXX-X.The valid format for an 11 digit value is XXXXX-XXXX-XX | Add | Optional | Type: Num. or Alphanum. Length: min-6, max-23* *defined by Issuing Agency structure. | NA | No | Yes
Previous DI
Issuing Agency | Name of Previous DI Issuing agency. | Autopopulated based on the Previous DI entered | None | Optional | NA | GS1; HIBCC; ICCBBA; NDC/NHRIC | No | Yes
Previous DI Number | Device Identifier assigned to a version/model of a medical device prior to the assignment of a new/substitute device identifier to that same version or model, without any changes to the device itself, due to merger and acquisition activity. | Enter Previous DI Number. The Previous DI must exist as a published Primary DI in another GUDID device record. Data type and field length are determined by the individual Issuing Agency structure. GS1: Num., DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanum., with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters | Add Edit | Optional | Type: Num. or Alphanum. Length: min-6, max-23* *defined by Issuing Agency structure. | NA | No | Yes
Package DI | Every device package shall bear a UDI, 21 CFR 801.20(a)(2). Package DIs do not need their own DI record; instead package information should be entered in the Package DI section of the Primary DI record for that device. According to 21 CFR 801.3, a package is defined as a fixed quantity of a particular version or model of a device.
Package DI Number | A device identifier for the package configuration that contains multiple units of the base package (does not include shipping containers). | Enter Package DI Number. Must be from same Issuing Agency as Primary DI. Data type and field length are determined by the individual Issuing Agency structure. GS1: Num., DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanum., with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters Examples: Box of Gloves = DI 101 4 Boxes of Gloves (DI 101) in a Carton = Package DI 201 (the UDI on the Carton) 5 Cartons (Pkg DI 201) in a Case = Package DI 301 (the UDI on the Case) 10 Boxes of Gloves (DI 101) in a Carton = Package DI 202 (the UDI on the Carton). | Add | Conditionally Required* *If device is available in higher levels of packaging | Type: Num. or Alphanum. Length: min-6, max-23* *defined by Issuing Agency structure. | NA | No | Yes
Quantity per Package | The number of packages with the same Primary DI or Package DI within a given packaging configuration. | Enter the number of devices per package. The quantity of a package configuration must be >1. Examples: Package – Carton, Pkg DI 201 contains 4 boxes of DI 101; the quantity per package is 4. Package – Case, Pkg DI 301 contains 5 cartons of Pkg DI 201; the quantity per package is 5. Package – Carton, Pkg DI 202 contains 10 boxes of DI 101; the quantity per package is 10. | Add | Conditionally Required* *If Package DI is entered | Type: Num. Length: 9 | NA | No | Yes
Contains DI Package | The Primary DI for the base package or the Package DI for any lower level package configuration contained within a given package configuration. | Choose a value from the drop down LOV. Examples: Package DI 201 (Carton) contains base package DI 101. Package DI 202 (Carton) contains base package DI 101. Package DI 301 contains lower level Package DI 201 (Carton). | Add | Conditionally Required* *If Package DI is entered | NA | DI numbers; base package and all lower levels of packaging | No | Yes
Package Type | Text to describe the outer packaging of the product and enables users to understand higher level packaging configurations. | Enter name or description of package. This field is free text. There is no implied definition or standard quantity to any package name. | Add | Optional | Type: Alphanum. Length: 20 | NA | No | Yes
Package Discontinue Date | Indicates the date this particular package configuration is discontinued by the labeler. | Choose date from calendar or manually enter in format yyyy-mm-dd. Discontinuation of a package is directly related to the discontinuation of the primary DI of the base package. However, a package can also be discontinued without the discontinuation of the base package. | Add | Conditionally Required* *If Package DI Number and Commercial Distribution End Date are entered, must also enter Package Discontinue Date | Type: Num. (date format) Length: 10 | NA | No | Yes
Package Status | Indicates whether the package is in commercial distribution as defined under 21 CFR 807.3(b). | Auto populated based on Package Discontinue Date. If Package DI and related elements are entered and no Package Distribution End Date is entered, the status is 'In Commercial Distribution.' | NA | Auto Populated | NA | In Commercial Distribution; Not in Commercial Distribution | No | Yes
Customer Contact
Customer Contact Phone | Phone number for the Customer contact; to be used by patients and consumers for device-related questions. | Enter phone number. For North American numbers, type 10-digit number with or without punctuation. For international numbers, start with "+" and type number without punctuation. This phone number could be the 1-800 number that appears on the device labeling or the company website. Labelers can identify a Customer Contact phone number and Customer Contact email address for each device record. If an email is entered and you don't have a Customer Contact phone number, please enter '9999999999' | Add Delete Edit | Conditionally Required* *ONLY required if Customer Contact Email is entered | Type: Num. Length: 10 (North American numbers); 20 (all others) | NA | No | Yes
Customer Contact Email | Email for the Customer contact; to be used by patients and consumers for device-related questions. | Enter email address. This email address could be the same one that appears on the device labeling or the company website. Labelers can identify a Customer Contact email and a Customer Contact phone number for each device record. If a phone number is entered and you don't have a Customer Contact email, please enter 'xx@xx.xx' | Add Delete Edit | Conditionally Required* *ONLY required if Customer Contact Phone is entered | Type: Alphanum. Length: 100 | NA | No | Yes
Device Status
Human Cell, Tissue or Cellular or Tissue-Based Product (HCT/P) | Indicates that the product contains or consists of human cells or tissues that are intended for implantation, transplantation, infusion, or transfer into a human recipient as defined under 21 CFR 1271.3. | Select checkbox if DI record is for a product defined under 21 CFR 1271.3 If checked, the labeler must assign and label each HCT/P device with a distinct identification code (DIC), per 21 CFR 1271.290(c). If checked, please ensure that you have a DIC code. The DIC may take the form of a Donation Identification Number (DIN) , serial number, lot number, or a combination of these production identifiers (PIs). IF Issuing Agency = ICCBBA, DIC = DIN Else IF Issuing Agency = GS1 or HIBCC then DIC = Serial Number OR DIC = Lot Number OR DIC = Serial + Lot Number Labelers of HCT/Ps regulated as medical devices should select the appropriate type of PI that appears on the label of the device. | Add Delete Edit | Optional If no data is provided, 'No' is stored | Type: Boolean | NA | No | Yes
Kit | Indicates that the device is a convenience, combination, in vitro diagnostic (IVD), or medical procedure kit. Kits are a collection of products, including medical devices, that are packaged together to achieve a common intended use and are being distributed as a medical device. | Select checkbox if DI record is for a kit. Do not check if the device is a constituent part of a kit. | None | Optional If no data is provided, 'No' is stored | Type: Boolean | NA | Yes | Yes
Combination Product | Indicates that the product is comprised of two or more regulated products that are physically, chemically, or otherwise combined or mixed and produced as a single entity; packaged together as a single package; or packaged separately for the intended use together as defined under 21 CFR 3.2(e). At least one of the products in the combination product must be a device in this case. | Select checkbox if DI record is for a combination product. Do not check if the device is a constituent part of a combination product. | None | Optional If no data is provided, 'No' is stored | Type: Boolean | NA | Yes | Yes
Premarket
Device Exempt from Premarket Submission | FDA Premarket submission is not required for this device. | Select checkbox if FDA has by regulation exempted this device from premarket submission requirements; or for pre-amendment devices that are not subject to premarket submission requirements. If left unselected, a 'No' is stored and a Premarket Submission Number should be entered below. | None | Conditionally Required* *Premarket Submission Number OR exempt status fulfills regulatory requirement | Type: Boolean | NA | No | Yes
FDA Premarket Submission Number | Number associated with the regulatory decision regarding the applicant’s legal right to market a medical device for the Submission types are: 510(k), De Novo, PMA, PDP, HDE, BLA, and NDA. | Enter current FDA Premarket Submission Number(s). Each DI record represents a version or model of a device. For each DI record, you must submit the premarket authorization number (if the model or version was originally approved in a PMA, HDE, BLA, or NDA supplement) and the supplement number through which you originally obtained approval or clearance or were granted classification for the version or model identified in the DI record, as required by 830.310(b)(11). FDA Premarket Numbers should be verified with the corresponding FDA premarket database to make sure the Number is consistent with the device record. Device records should be updated with additional numbers in the future, as needed. FDA Premarket Supplement number is collected as a separate field. Example: De Novo #123456 should be entered as 'DEN123456.' Example: PMA # P190010/S46 should be entered as 'P190010' in the 'FDA Premarket Submission Number' field and '46' in the 'FDA Premarket Supplement Number' field. | Add | Conditionally Required* *Premarket Submission Number OR exempt status fulfills regulatory requirement. | Type: Alphanum. Length: 8 | NA | No | No We are planning to release this data element in the near future.
Supplement Number | Number assigned by FDA to a supplemental application for approval of a change in a medical device with an approved PMA, HDE, or PDP. | Enter all valid Supplement Numbers. Each DI record represents a version or model of a device. For each DI record, you must submit the premarket authorization number (if the model or version was originally approved in a PMA, HDE, BLA, or NDA supplement) and the supplement number through which you originally obtained approval for the version or model identified in that DI record, as required by 830.310(b)(11). Although not all supplements are applicable to a given model or version, if FDA approves a subsequent supplement applicable to that version or model, the GUDID DI record must be updated with that supplement number, in accordance with 21 CFR 830.330(b). 30 day notice supplements should be submitted ONLY if the 30 day notice impacts the device design specifications, or performance of the finished devices. Do not enter alpha characters. Example: Supplement 4 should be entered as '4'. | Add | Conditionally Required* *Premarket Submission Number OR exempt status fulfills regulatory requirement. | Type: Num. Length: CDRH - min-1, max-3 CBER - min-1, max-4 | NA | No | No We are planning to release this data element in the near future.
FDA Product Code
Product Code | Categorization for devices used by the FDA. | Enter all applicable Product Codes, three-letter code. For all PMA, 510k, HDE and De Novo devices, product codes are assigned in the FDA approval, clearance or classification letter, respectively. If the product code is down classified then use the same product code that was used to register and list the device. For most De Novo devices, the product code can be found in the reclassification order. | Add Delete Edit | Conditionally Required* *Unless device is a kit or IVD with a BL premarket submission number | Type: Alpha Length: 3 | FDA Product Code list | No | Yes
Product Code Name | Name associated with the three-letter Product Code. | Auto populated based on 3-letter Product Code | NA | Auto Populated | NA | NA | No | Yes
FDA Listing
FDA Listing Number | Number assigned by FDA during Registration and Listing to all devices in commercial distribution, regardless of pre-market authorization requirements per 21 CFR 807.28(f). | Enter all relevant listing numbers that enable the labeler to commercially distribute the given version or model of device. Listing number is optional for HCT/P devices with a BLA premarket number. | Add | Conditionally Required* *Unless device is an HCT/P with a BL premarket submission number | Type: Alphanum. Length: 7 | NA | No | No
GMDN
Code | GMDN Preferred Term (PT) Code is a unique five-digit code used to identify common device types. This PT Code is assigned to medical devices and related health care products for the purposes of grouping and categorization. | Enter all applicable GMDN Preferred Term Codes or FDA PT Codes. Each device record must have at least one assigned GMDN Code/FDA PT Code; DI records are allowed >1 GMDN Code/FDA PT Code, if necessary. Must enter GMDN Code OR FDA PT Code, please don't enter both codes for the same GMDN Name and Definition. For GMDN Codes: Enter only the 5-digit number, omit the 'P' For FDA PT Codes: Enter the 4-letter code. The FDA PT Codes are assigned to each GMDN term, used in place of a GMDN Code. They enable labelers to assign a GMDN term to their GUDID submission until a GMDN Code can be obtained from the GMDN Agency. The FDA PT Codes can be found in the Find FDA PT Code Module on the GUDID website. For more information, see the GUDID Final Guidance. The FDA PT Codes can be found in the Find FDA PT Code Module on the GUDID website. | Add Delete Edit | Required | GMDN -- Type: Num. Length: 5 FDA PT Code: Type: Alpha Length: 4 | NA | No | No
Name | Name of the common device type associated with the GMDN Preferred Term Code/FDA PT Code. | Auto populated based on GMDN Preferred Term Code/FDA PT Code. | NA | Auto Populated | NA | NA | No | Yes
Definition | Definition of the common device type associated with the GMDN Preferred Term Code/FDA PT Code. | Auto populated based on GMDN Preferred Term Code/FDA PT Code. | NA | Auto Populated | NA | NA | No | Yes
Device Characteristics
For Single-Use | Indicates that the device is intended for one use or on a single patient during a single procedure. | Choose Yes/No from the drop down list. | None | Required | Type: Boolean | Yes/No | Yes | Yes
Production Identifier(s) in UDI
Lot or Batch Number | The number assigned to one or more device(s) that consist of a single type, model, class, size, composition, or software version that are manufactured under essentially the same conditions and that are intended to have uniform characteristics and quality within specified limits. This number is required to be part of the UDI when included on the label in order to provide the means to track the device back to its manufacturing source or otherwise allow the history of the device manufacturing, packaging, labeling, distribution and use to be determined. | Choose Yes/No from the drop down list. For stand-alone software, select Yes to indicate that the software version number will be represented as a Lot or Batch number | Add Delete Edit | Required | Type: Boolean | Yes/No | No | Yes
Manufacturing Date | The date on which a device is manufactured. This date is required to be part of the UDI when included on the label in order to provide the means to track the device back to its manufacturing source or otherwise allow the history of the device manufacturing, packaging, labeling, distribution and use to be determined. | Choose Yes/No from the drop down list. | Add Delete Edit | Required | Type: Boolean | Yes/No | No | Yes
Serial Number | The number that allows for the identification of a device, indicating its position within a series. This number is required to be part of the UDI when included on the label in order to provide the means to track the device back to its manufacturing source or otherwise allow the history of the device manufacturing, packaging, labeling, distribution and use to be determined. | Choose Yes/No from the drop down list. | Add Delete Edit | Required | Type: Boolean | Yes/No | No | Yes
Expiration Date | The date by which the label of a device states the device must or should be used. This date is required to be part of the UDI when included on the label in order to provide the means to track the device back to its manufacturing source or otherwise allow the history of the device manufacturing, packaging, labeling, distribution and use to be determined. | Choose Yes/No from the drop down list. | Add Delete Edit | Required | Type: Boolean | Yes/No | No | Yes
Donation Identification Number | The Donation Identification Number (DIN) is applicable to devices that are also regulated as HCT/Ps and is a number that is assigned to each donation. This number/code is required to be part of the UDI when included on the label in order to provide the means to track the device back to its manufacturing source or otherwise allow the history of the device manufacturing, packaging, labeling, distribution and use to be determined. | Choose Yes/No from the drop down list. This PI is only applicable to HCT/P products regulated as medical devices. IF Issuing Agency = ICCBBA, DIC = DIN Else IF Issuing Agency = GS1 or HIBCC then DIC = Serial Number OR DIC = Lot Number OR DIC = Serial + Lot Number | Add Delete Edit | Required | Type: Boolean | Yes/No | No | Yes
Latex Information
Device required to be labeled as containing natural rubber latex or dry natural rubber (21 CFR 801.437). | Indicates that the device or packaging contains natural rubber that contacts humans as described under 21 CFR 801.437. Choosing 'Yes' indicates that the device label or packaging contains one of the following statements: (1) "Caution: This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions", (2) This Product Contains Dry Natural Rubber", (3) Caution: The Packaging of This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions" or (4) "The Packaging of This Product Contains Dry Natural Rubber". | Choose Yes/No from the drop down list. | None | Required | Type: Boolean | Yes/No | Yes | Yes
Device labeled as "Not made with natural rubber latex" | Indicates that natural rubber latex was not used as materials in the manufacture of the medical product and container and the device labeling contains this information. Only applicable to devices not subject to the requirements under 21 CFR 801.437. Not all medical products that are NOT made with natural rubber latex will be marked. | Select checkbox if appropriate. Only applicable if the response to "Device required to be labeled as containing natural rubber latex or dry natural rubber" is "No". Optional element for labelers who include a statement of 'latex-free' on their label or in their labeling. FDA finds these statements: 'latex-free' and 'does not contain latex', to be not scientifically supportable and strongly recommends they not be used in medical product labeling. Instead FDA recommends the use of the statement 'Not made with natural rubber latex." It is not assumed that all devices NOT made with natural rubber latex are marked; therefore this is an optional element for the labelers who choose to make a statement in the labeling. | Add Delete Edit | Optional If no data is provided, "No" is stored | Type: Boolean | NA | No | Public
Prescription Status
Prescription Use (Rx) | Indicates that the device requires a prescription to use. | Select checkbox if appropriate. Can select both Rx and OTC for one DI record. | Add Delete Edit | Optional If no data is provided, "No" is stored | Type: Boolean | NA | No | Yes
Over the Counter (OTC) | Indicates that the device does not require a prescription to use and can be purchased over the counter (OTC). | Select checkbox if appropriate. Can select both Rx and OTC for one DI record. | Add Delete Edit | Optional If no data is provided, "No" is stored | Type: Boolean | NA | No | Yes
MRI Safety Status
What MRI safety information does the labeling contain? | Indicates the MRI Safety Information, if any, that is present in the device labeling. Please see the ASTM F2503-13 standard for more information. | Choose a value from the drop down LOV. The final rule does not require MRI-compatibility testing; it only requires submission of information regarding MRI-compatibility that the labeler already possesses. | Edit* *ONLY if changing from 'Labeling does not contain…' to other MR status (Safe, Unsafe, Conditional). Otherwise, NO changes are allowed. | Required | NA | Refer to List of Values (LOV) sheet | Yes | Yes
Clinically Relevant Size
Size Type | Dimension type for the clinically relevant measurement of the medical device. | Choose a value from the drop down LOV. If you do not see your size type in the List of Values (LOV), then submit your suggestion for review as a new LOV via FDA UDI Help Desk. Submit with a subject line List of Values. Include the GMDN term(s) that the LOV is pertinent to. If the desired Size Type is not in the current list, select 'Size Text, specify' and the data element 'Size Type Text' will appear (see below). It is expected that the 'Size Text, specify' will only be available for a limited time. Use this option to help us build a list of values that are appropriate for your device type. GUDID reserves the right to review all suggestions before adding values to the Size Type LOV. More than one Size Value per Type and more than one Size Type may be added to each DI record. Entry Example - 8cm by 2 cm - requires 2 entries Entry 1 - Size Type = Length, Size Value = 8, Size Unit of Measure = Centimeter Entry 2 - Size Type = Width, Size Value = 2, Size Unit of Measure = Centimeter | Add | Conditionally Required* *If device is available in more than one size | NA | Refer to List of Values (LOV) sheet | No | Yes
Size Value | Numeric value for the clinically relevant size measurement of the medical device. | Enter numeric value for size. Decimals are accepted; fractions are not accepted. Each Size Value should be entered separately. GUDID is not accepting Size Value as a range at this time. Entry Example - 8cm by 2 cm - requires 2 entries Entry 1 - Size Type = Length, Size Value = 8, Size Unit of Measure = Centimeter Entry 2 - Size Type = Width, Size Value = 2, Size Unit of Measure = Centimeter | Add | Conditionally Required* *Required if device is available in more than one size | Type: Num. Length: 40 | NA | No | Yes
Size Unit of Measure | The unit of measure associated with each clinically relevant size. | Choose a value from the drop down. Entry Example - 8cm by 2 cm - requires 2 entries Entry 1 - Size Type = Length, Size Value = 8, Size Unit of Measure = Centimeter Entry 2 - Size Type = Width, Size Value = 2, Size Unit of Measure = Centimeter | Add | Conditionally Required* *Required if device is available in more than one size | NA | Refer to List of Values (LOV) sheet | No | Yes
Size Type Text | Additional undefined device size not represented in the GUDID Size Type LOV. | Enter Size Type, Size Unit and Unit of Measure for each entry. | Add | Conditionally Required* *Required if 'Size Text, specify' is selected above | Type: Alphanum. Length: 200 | NA | No | Yes
Storage and Handling
Storage and Handling Type | Indicates storage and handling requirements that are required for the device including temperature, humidity, and atmospheric pressure. | Choose a value from the drop down LOV. Conditions of the Storage and Handling Type are measured below as a range, with a Low Value and a High Value. More than one Storage and Handling Type can be added per device record. | Add Delete Edit | Optional | NA | Refer to List of Values (LOV) sheet | No | Yes
Low Value | Indicates the low value for storage and handling requirements. | Enter a number for Low Value. Must enter at least one value, Low or High but can enter both Low Value and High Value, if needed. When Storage and Handling value is a range, this is the lower end of that range When Storage and Handling value is not to fall below a minimum value, enter the number here When Storage and Handling value is exactly a value, enter the value here and in Storage and Handling High Value | Add Delete Edit | Conditionally Required* *One value (Low or High) is required if Storage and Handling Type is added to the device record | Type: Num. Length: 6 | NA | No | Yes
High Value | Indicates the high value for storage and handling requirements. | Enter a number for High Value. Must enter at least one value, Low or High but can enter both Low Value and High Value, if needed. When Storage and Handling value is a range, this is the higher end of that range When Storage and Handling value is not to rise above a maximum value, enter the number here When Storage and Handling value is exactly a value, enter the value here and in Storage and Handling Low Value | Add Delete Edit | Conditionally Required* *One value (Low or High) is required if Storage and Handling Type is added to the device record | Type: Num. Length: 6 | NA | No | Yes
Unit of Measure | The unit of measure associated with the storage and handling conditions. | Choose a value from the drop down LOV. | Add Delete Edit | Conditionally Required* *Required if Storage and Handling Type is added to the device record | NA | Refer to List of Values (LOV) sheet | No | Yes
Special Storage Conditions | Indicates any special storage requirements for the device. | Enter any other storage conditions. For devices kept at room temperature, or other standard conditions, input that information here. | Add Delete Edit | Conditionally Required* *Required if 'Special Storage Conditions' is selected above | Type: Alphanum. Length: 200 | NA | No | Yes
Sterilization Method
Device Packaged as Sterile | Indicates the medical device is free from viable microorganisms. See ISO/TS 11139. | Choose Yes/No from the drop down list. The two Sterilization Method questions are independent of each other; this element is designed to capture information about the device as it enters Commercial Distribution. These data elements are not designed to capture sterilization procedures executed by the manufacturer or labeler. | None | Required | Type: Boolean | Yes/No | Yes | Yes
Requires Sterilization Prior to Use | Indicates that the device requires sterilization prior to use. | Choose Yes/No from the drop down list. The two Sterilization Method questions are independent of each other; this element is designed to capture information about the device before it can safely encounter a patient, regardless of whether the device is single use or reused after reprocessing. These data elements are not designed to capture sterilization procedures executed by the manufacturer or labeler. If answered 'Yes', at least one Sterilization Method (below) must be selected. | None | Required | Type: Boolean | Yes/No | Yes | Yes
Sterilization Method | Indicates the method(s) of sterilization that can be used for this device prior to device use on the patient. | Choose a value from the drop down LOV. Only applicable if the answer to 'Requires Sterilization Prior to Use' is 'Yes'; otherwise, the LOV will remain inactive. Note that the sterilization method must be applied by the user prior to device use on patient. The Entry LOVs represent the sterilization methods recognized by the CDRH Infection Control Branch. Methods selected should be only those approved for each device by the CDRH Office of Device Evaluation. | Add Delete Edit | Conditionally Required* *if 'Requires Sterilization Prior to Use' is marked 'Yes' | NA | Refer to List of Values (LOV) sheet | No | Yes
Public versioning
Public Device Record Key | System generated unique key to identify a device record, even if the Primary DI changes | Auto populated by GUDID system | NA | Auto Populated | Type: Alphanum. | NA | No | Yes
Public Version Number | System generated number to track public release updates to a device record | Auto populated by GUDID system | NA | Auto Populated | Type: Num. |  | No | Yes
Public Version Date | Public Release Date of the device record (new/update) | Auto populated by GUDID system | NA | Auto Populated | Type: Date Format: (yyyy-mm-dd HH:MM) | NA | No | Yes
Public Version Status | System generated status of a device record associated with a record version change. | Auto populated by GUDID system | NA | Auto Populated | Type: Alphanum. | New; Update; Delete | No | Yes
Footnotes
1 | Add = Addition of new data is allowed; Delete = Deletion of entered data is allowed; Edit = Editing of entered data is allowed; None = NO edit, add, or delete are allowed; NA = data element is not able to be changed directly; most are ‘auto-populated’ fields whose information depends on another data element Note: The above do not apply if the device is “unlocked” for editing. For more information on “unlocking” device records for editing, please visit www.fda.gov/udi
2 | See 21 CFR 830.310 and 830.340 for required data elements.
3 | Alpha = alpha only; Num. = numeric; Alphanum. = alphanumeric; Date Format = enter as yyyy-mm-dd; NA = data elements with a checkbox or LOV that don’t require text entry
4 | Most of the information presented here is applicable to GUDID HL7 SPL submissions, but there are some differences pertinent to each submission option. Please refer to the HL7 SPL Implementation package of files for additional details on HL7 SPL xml file submission option.
5 | Acronyms Used - See Acronyms sheet
6 | Data elements required in the implantable device list as part of electronic health records per the ONC 2015 edition Health IT certification criteria regulation 45 CFR 170.315(a)(14). The data elements are primary DI, brand name, model/version, company name, MR safety, Latex, Description(GMDN/SNOMED), and Production Identifiers (Lot, serial number, expiration date, manufacture date, and DIC (see Donation Identification Number data element). See the Helpful Links sheet for link to the regulation.

## ChangeLog
GUDID Data Elements Reference Table
Change Log
Note: All changes appear in order with the most recent date shown first. For previous changes, scroll to the bottom of the list.
Date | Data Sheet | Data Element | Change
2018-08-15 | GUDID Data Elements | Brand Name | Updated Data Entry Note to read - Enter the brand name specific to this version/model of the device. You may include the ® and/or TM symbol, if applicable. If there is no brand name for your device then enter NA or the name commonly used to refer to the device or family of devices. Information entered in this field should be greater than one character in length, and a name that refers to the device must be entered.
| Helpful Links |  | Modify link to OpenFDA
2018-05-11 | GUDID Data Elements |  | Add header note to read - According to 830.330(b), The labeler of a device shall submit to FDA an update to the information required by § 830.310 whenever the information changes. The updated information must be submitted no later than the date a device is first labeled with the changed information. If the information does not appear on the label of a device, the updated information must be submitted within 10 business days of the change.
| General Notes |  | General Notes sheet was added
| Acronyms |  | Acronyms sheet was added
| List of Values (LOV) |  | List of Values (LOV) sheet was added
| Sample Label |  | Sample Label sheet was added
| Helpful Links |  | Helpful Links sheet was added
| GUDID Data Elements | Entry List of Values column | Data elements with a long List of Values (LOV) within this cell have been moved to a separate sheet called List of Values (LOV).
| GUDID Data Elements | Footnote 5 | Acronyms list has been expanded and listed in a new Acronyms sheet
| GUDID Data Elements | Footnote 6 | Added to Data Elements heading as part of ONC regulation. Data elements required in the implantable device list as part of electronic health records per the ONC 2015 edition Health IT certification criteria regulation 45 CFR 170.315(a)(14). The data elements are primary DI, brand name, model/version, company name, MR safety, Latex, Description(GMDN/SNOMED), and Production Identifiers (Lot, serial number, expiration date, manufacture date, and DIC (see Donation Identification Number data element). https://www.healthit.gov/sites/default/files/2015Ed_CCG_a14-Implantable-device-list.pdf
| GUDID Data Elements | Primary DI | Updated Description to read - The Primary Device Identifier (DI) is the DI portion of the UDI placed on the lowest package level of a device that is required to meet UDI label requirements. If the device is not packaged, the UDI may be on the device itself, thereby satisfying both the UDI label and the direct mark (DM) requirement if the UDI is intended to be permanent. The primary DI is the main (primary) lookup for a medical device and meets the requirements to uniquely identify a device through its distribution and use.
| GUDID Data Elements | Primary DI | Updated Data Entry Note to read - Enter the Device Identifier (DI) Number. If using GS1 as an issuing agency, ensure the check digit is correct and valid as per the GS1 guidelines. If using HIBCC or ICCBBA as an issuing agency, do not include the check digit/character as part of your DI. Data type and field length are determined by the individual Issuing Agency structure. IF Device IS Packaged Primary DI = DI portion of the UDI on the lowest package level OR IF Device is NOT Packaged Primary DI = DI on the device itself, could be DM DI GS1: Numeric (Num.), DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanumeric (Alphanum.), with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters
| GUDID Data Elements | Labeler DUNS number | Change from No to Yes in Publicly Released column
| GUDID Data Elements | Brand Name | Updated Description to read - A brand name is a trade/proprietary name assigned by the device labeler, and under which the device is sold, distinguished from other similar devices, and recognized by the user or purchaser. A brand name is often registered and/or protected by a trademark.
| GUDID Data Elements | Brand Name | Updated Data Entry Note to read - Enter the brand name specific to this version/model of the device. You may include the ® and/or TM symbol, if applicable. If there is no brand name for your device then enter the name commonly used to refer to the device or family of devices. Information entered in this field should be greater than one character in length, and a name that refers to the device must be entered.
| GUDID Data Elements | Version or Model | Updated Data Entry Note to read - Enter the Version or Model. Version/Model can be any distinguishing string of letters and/or numbers. If the device does not have a version or model, enter a concept that can be used to identify all devices that have specifications, performance, size, and composition within limits set by the labeler.
| GUDID Data Elements | Catalog Number | Updated Description to read - The catalog, reference, reorder or product number used by the labeler for business and clinical transactions to identify a particular product; this number may be found on the device label or labeling.
| GUDID Data Elements | Catalog Number | Updated Data Entry Note to read - The catalog number is critical for matching UDI to legacy device identification data used in the item masters, inventory systems, EHRs, and registries. The DI of the UDI will replace the catalog number for recalls and adverse event reports in the future. To populate this field: Use the catalog, reference, reorder or product number on the device label. If there is no catalog, reference, reorder or product number use the number reported in a business/clinical transaction or that would be reported in case of a recall or adverse event.
| GUDID Data Elements | Catalog Number | Updated Required in Database column to read - Recommended - Critical for identification and adoption
| GUDID Data Elements | Device Description | Updated Description to read - Additional information about the device that is not already captured as a distinct GUDID data element.
| GUDID Data Elements | Device Description | Updated Data Entry Note to read - Device description should include any description found on the device label, package insert, or approved or cleared indications for use to support user identification. If you have multiple device records with similar information, this field should help users differentiate among the device records. You may include URL links to: the label, instructions for use, device pictures/schematics. The inclusion of a URL is to supplement the description of the device. URLs can change over time. If a URL is included, please keep the URL up to date. Duplicating information already provided as distinct data elements in GUDID, e.g. Brand name, company name, etc. is not necessary. •Note: This field appears in the results summary from searches in AccessGUDID.
| GUDID Data Elements | Commercial Distribution End Date | Choose date from calendar or manually enter date in format yyyy-mm-dd. Do not populate this field for deactivating/deleting records. Please contact the FDA UDI Help Desk at FDA UDI Help Desk. If you do not know an exact commercial distribution end date, please do not populate an estimated date.
| GUDID Data Elements | DM DI Different from Primary DI | Updated Data Entry Note to read - IF Device IS Packaged AND DM DI ≠ Primary DI on device label THEN Select checkbox DM DI Different from Primary DI and Enter DM DI in DM DI field
| GUDID Data Elements | DM DI Number | Updated Description to read - An identifier that is permanently marked directly on the medical device. Certain devices are required to be directly marked. Under 21 CFR 801.45(b), the direct mark DI may be identical to or different from the Primary DI.
| GUDID Data Elements | DM DI Number | Updated Data Entry Note to read - If the direct mark is identical to the UDI that appears on the device label, do not enter a direct mark DI. If the direct mark is different from the UDI that appears on the device label, the direct mark DI is the DI portion of the direct mark. If a device is not packaged and the UDI is on the device itself and intended to be permanent, direct mark requirements are satisfied, enter the direct mark DI as the primary DI and do not enter a DM DI. The above text is also applicable to devices that are not subject to 21 CFR 801.45, but are voluntarily direct marked with a UDI by the labeler. IF Device IS Packaged AND DM DI = Primary DI on device label THEN Do not enter DM DI ELSE IF DM DI ≠ Primary DI on device label THEN Enter DM DI IF Device IS NOT Packaged AND DI permanently marked on the device (i.e., DM DI) AND DM DI = Primary DI THEN Do not enter DM DI; DM DI is the Primary DI GS1: Num., DI may be 12, 13 or 14 digit numeric value. System will append zeros to the left or beginning of DI to save 14 digit value, include the check digit as part of the DI HIBCC: Alphanum., with 6-23 characters, DO NOT include the check digit as part of the DI ICCBBA: Alphanum., with 10 or 16 characters
| GUDID Data Elements | Human Cell, Tissue or Cellular or Tissue-Based Product (HCT/P) | Updated Data Entry Note to read - Select checkbox if DI record is for a product defined under 21 CFR 1271.3 If checked, the labeler must assign and label each HCT/P device with a distinct identification code (DIC), per 21 CFR 1271.290(c). If checked, please ensure that you have a DIC code. The DIC may take the form of a Donation Identification Number (DIN) , serial number, lot number, or a combination of these production identifiers (PIs). IF Issuing Agency = ICCBBA, DIC = DIN Else IF Issuing Agency = GS1 or HIBCC then DIC = Serial Number OR DIC = Lot Number OR DIC = Serial + Lot Number Labelers of HCT/Ps regulated as medical devices should select the appropriate type of PI that appears on the label of the device.
| GUDID Data Elements | FDA Premarket Submission Number | Updated Publicly Released column to read - No; We are planning to release this data element in the near future.
| GUDID Data Elements | FDA Premarket Submission Number | Updated Data Entry Note to read - Enter current FDA Premarket Submission Number(s). Each DI record represents a version or model of a device. For each DI record, you must submit the premarket authorization number (if the model or version was originally approved in a PMA, HDE, BLA, or NDA supplement) and the supplement number through which you originally obtained approval or clearance or were granted classification for the version or model identified in the DI record, as required by 830.310(b)(11). FDA Premarket Numbers should be verified with the corresponding FDA premarket database to make sure the Number is consistent with the device record. Device records should be updated with additional numbers in the future, as needed. FDA Premarket Supplement number is collected as a separate field. Example: De Novo #123456 should be entered as 'DEN123456.' Example: PMA # P190010/S46 should be entered as 'P190010' in the 'FDA Premarket Submission Number' field and '46' in the 'FDA Premarket Supplement Number' field.
| GUDID Data Elements | Product Code | Updated Description to read - Categorization for devices used by the FDA.
| GUDID Data Elements | Product Code | Updated Data Entry Note to read - Enter all applicable Product Codes, three-letter code. For all PMA, 510k, HDE and De Novo devices, product codes are assigned in the FDA approval, clearance or classification letter, respectively. If the product code is down classified then use the same product code that was used to register and list the device. For most De Novo devices, the product code can be found in the reclassification order.
| GUDID Data Elements | Donation Identification Number | Updated Data Entry Note to read - Choose Yes/No from the drop down list. This PI is only applicable to HCT/P products regulated as medical devices. IF Issuing Agency = ICCBBA, DIC = DIN Else IF Issuing Agency = GS1 or HIBCC then DIC = Serial Number OR DIC = Lot Number OR DIC = Serial + Lot Number
| GUDID Data Elements | What MRI safety information does the labeling contain? | Updated Entry List of Values (LOV) to read - Refer to List of Values (LOV) sheet
| GUDID Data Elements | Size Type | Updated Data Entry Note to read - Choose a value from the drop down LOV. If you do not see your size type in the List of Values (LOV), then submit your suggestion for review as a new LOV via FDA UDI Help Desk. Submit with a subject line List of Values. Include the GMDN term(s) that the LOV is pertinent to. If the desired Size Type is not in the current list, select 'Size Text, specify' and the data element 'Size Type Text' will appear (see below). It is expected that the 'Size Text, specify' will only be available for a limited time. Use this option to help us build a list of values that are appropriate for your device type. GUDID reserves the right to review all suggestions before adding values to the Size Type LOV. More than one Size Value per Type and more than one Size Type may be added to each DI record. Entry Example - 8cm by 2 cm - requires 2 entries Entry 1 - Size Type = Length, Size Value = 8, Size Unit of Measure = Centimeter Entry 2 - Size Type = Width, Size Value = 2, Size Unit of Measure = Centimeter
| GUDID Data Elements | Size Type | Updated Entry List of Values (LOV) to read - Refer to List of Values (LOV) sheet
| GUDID Data Elements | Size Unit of Measure | Updated Entry List of Values (LOV) to read - Refer to List of Values (LOV) sheet
| GUDID Data Elements | Storage and Handling Type | Updated Entry List of Values (LOV) to read - Refer to List of Values (LOV) sheet
| GUDID Data Elements | Storage and Handling Unit of Measure | Updated Entry List of Values (LOV) to read - Refer to List of Values (LOV) sheet
| GUDID Data Elements | Storage and Handling High Value | Updated Data Entry Note to read - Enter a number for High Value. Must enter at least one value, Low or High but can enter both Low Value and High Value, if needed. When Storage and Handling value is a range, this is the higher end of that range When Storage and Handling value is not to rise above a maximum value, enter the number here When Storage and Handling value is exactly a value, enter the value here and in Storage and Handling Low Value
| GUDID Data Elements | Storage and Handling Low Value | Updated Data Entry Note to read - Enter a number for Low Value. Must enter at least one value, Low or High but can enter both Low Value and High Value, if needed. When Storage and Handling value is a range, this is the lower end of that range When Storage and Handling value is not to fall below a minimum value, enter the number here When Storage and Handling value is exactly a value, enter the value here and in Storage and Handling High Value
| GUDID Data Elements | Sterilization Method | Updated Entry List of Values (LOV) to read - Refer to List of Values (LOV) sheet
| GUDID Data Elements | Public Version Status | Added a new data element
2018-03-28 | GUDID Data Elements | Public Device Record Key | Added a new data element
|  | Public Version Number | Added a new data element
|  | Public Version Date | Added a new data element
|  | Footnote 1 | Added a note to footnote 1 -- Note: The above do not apply if the device is “unlocked” for editing. For more information on “unlocking” device records for editing, please visit www.fda.gov/udi
2017-03-27 | GUDID Data Elements | Primary DI | Updated Data Entry Notes to indicate when to include check digit as part of DI in GUDID
|  | Unit of Use DI | Slight tweaks to the definition to add more clarity
|  | Secondary DI | Changed "Edit Rules after Grace Period" column to read -- "Add" Secondary DIs may now be added to GUDID after the DI record grace period
|  | Catalog Number | Updated Data Entry Notes to read - Catalog Number is critical for UDI adoption in electronic health records. Please provide catalog number as part of your device record.
|  | Previous DI Issuing Agency | Added new row
|  | Previous DI | Added new row
|  | Unit of Use DI | Updated description to indicate that Unit of Use DI is an unmarked device identifier that does not appear on the label of the device
Added a new column titled "Publicly Released" -- indicates data element values that are released to the public on AccessGUDID and OpenFDA
2016-04-14 | GUDID Data Elements | Commercial Distribution End Date | Updated description to read as follows -- Indicates the date the device is no longer offered for commercial distribution by the labeler on record. See 21 CFR 807.3(b) for exceptions. The device may or may not still be available for purchase in the marketplace.
|  | Secondary DI Number | Added NDC/NHRIC as an Issuing Agency type along with data type and length
|  | Listing Number | Updated Data Entry Notes and Required Columns to show that Listing Number is optional for HCT/P devices with a BL premarket number.
|  | Sterilization Method | Updated description to read as follows -- Indicates the method(s) of sterilization that can be used for this device prior to device use on the patient. Updated Data Entry notes
2015-05-01 | GUDID Data Elements | Version or Model Number | Updated Length = 80
|  | Catalog | Updated Length = 80
|  | Commercial Distribution End Date | Updated Description to read as follows -- Indicates the date the device is no longer held or offered for sale by the labeler on record. See 21 CFR 807.3(b) for exceptions. The device may or may not still be available for purchase in the marketplace.
2015-01-12 | GUDID Data Elements | Unit of Use DI Number | Updated Data Entry Notes to indicate -- Must be from same Issuing Agency as Primary DI.
|  | DM DI Number | Updated Data Entry Notes to indicate -- Must be from same Issuing Agency as Primary DI.
|  | Package DI Number | Updated Data Entry Notes to indicate -- Must be from same Issuing Agency as Primary DI.
|  | Customer Contact Email | Updated Data Entry Notes to change email address from "xxx@xxx.xxx" to -- If a phone number is entered and you don't have a Customer Contact email, please enter 'xx@xx.xx'
|  | Premarket Submission Supplement Number | Updated Data Type & Length Column from Length = 4 digits to Length = CDRH 1-3 digits, CBER 1-4 digits
|  | GMDN Code | Added information about FDA PT Codes to Data Entry Notes
|  | Production Identifiers in UDI - Lot or Batch, Manufacturing Date, Serial Number, Expiration Date, Donation Identification Number | Updated description
|  | Size Unit of Measure | Updated Entry List of Values Column -- Added Millibar for pressure; Removed -- Microgram per Total Volume; Milligram per Total Volume
|  | Storage and Handling Low Value | Added to Data Entry Notes -- When Storage and Handling value is a range, this is the lower end of that range When Storage and Handling value is less than a value, enter the number here When Storage and Handling value is exactly a value, enter the value here and in Storage and Handling High Value
|  | Storage and Handling High Value | Added to Data Entry Notes -- When Storage and Handling value is a range, this is the higher end of that range When Storage and Handling value is greater than a value, enter the number here When Storage and Handling value is exactly a value, enter the value here and in Storage and Handling Low Value
|  | Storage and Handling Unit of Measure | Updated Entry List of Values Column -- Added Millibar for Storage and Handling Pressure
Why do we have Type/Limit in Data Entry Notes???

## General Notes
General Notes
Double check your data for accuracy and completeness.
When entering in alpha characters into an attribute, enter in English characters.
Do not enter in special characters like double quotes (") or single quotes ('). This has caused problems with data users who download and import this data.

## Acronyms
Acronyms
510(k) - Premarket Notification
AccessGUDID - public website for GUDID database
Alphanum.-Alphanumeric
ASTM - ASTM International formerly known as American Society Testing Materials
BL - Biologic Licensing
BLA - Biologic Licensing (Application)
C-Code - Concept Code
CDRH - FDA Centers for Devices and Radiologic Health
CFR- Code of Federal Regulations
D&B - Dun & Bradstreet
DEN - De Novo
DERT - Data Elements Reference Table
DI - Device Identifier
DIC - Distinct Identification Code
DM - Direct Mark
DIN - Donation Identification Number
DUNS - Data Universal Numbering System
EA - each
FDA - Food and Drug Administration
GMDN - Global Medical Device Nomenclature
GS1 - Global Standards One
GUDID - Global Unique Device Identification Database
HCT/P - Human Cell, Tissue or Cellular or Tissue-Based Product
HDE - Humanitarian Device Exemption
HIBCC - Health Industry Business Communications Council
HL7 - Health Level Seven
ICCBBA - International Council for Commonality in Blood Banking Automation
IMDRF - International Medical Device Regulators Forum
ISO - International Organization for Standardization
LOV - List of Values
MR - Magnetic Resonance
MRI - Magnetic Resonance Imaging
NA or N/A - Not Applicable
NCIt - National Cancer Institute's Thesaurus
NDA - New Drug Application
NDC - National Drug Code
NHRIC - National Health Related Items Code
Num.- Numeric
ONC - Office of National Coordinator
OTC - Over The Counter
PDP - Product Development Protocol
PI - Production Identifier
PMA - Pre Market Authorization
PT - GMDN Preferred Term
® - Registered trademark symbol
Rx - prescription
SNOMED - Systematized Nomenclature of Medicine
SPL - Structured Product Labeling
TM - Trademark
UCUM - Unified Code for Units of Measure
UDI - Unique Device Identifier
URL - Uniform Resource Location (internet shortcut)
USPTO - United States Patent and Trademark Office

## List of Values (LOV)
List of Values 1
List of Values |  | Units of Measure |  | FDA CDRH-GUDID definition
Data Element | C-Code 2 | Description | Code 3
Size Types
Angle | C67505 | Degree | deg | The inclination of one line to another or the plane of one object to another.
Area
Size Types |  | Description | Code
Area/Surface Area | C25244 | Square centimeter | cm2 | The extent of a 2-dimensional surface enclosed within a boundary.
|  | Square foot | [sft_i]
|  | Square inch | [sin_i]
|  | Square meter | m2
|  | Square millimeter | mm2
Gauge - French
Size Types |  | Description | Code
Catheter Gauge | C101680 | French | [Ch] | A number representing the outer diameter of a catheter where each integer represents 1/3 of a millimeter.
Gauge - Needle
Size Types |  | Description | Code
Needle Gauge | C101687 | Gauge | G | A number representing the outer diameter of a hypodermic needle.
Note: This unit of measure is not provided in UCUM. The code value is one specific to the GUDID implementation
Length Size
Size Types |  | Description | Code
Circumference | C64265 | Centimeter | cm | The length of the closed curve of a circle; the size of something as given by the distance around it.
Depth | C25333 | Decimeter | dm | The extent downward or inward; the perpendicular measurement from the surface downward to determine deepness.
Height | C25347 | Feet | [ft_i] | The vertical measurement or distance from the base to the top of an object; the vertical dimension of extension.
Length | C25334 | Femtometer | fm | The linear extent in space from one end of something to the other end, or the extent of something from beginning to end.
Lumen/Inner Diameter | C101685 | Inch | [in_i] | The length of a straight line passing through the center of the inner open space or cavity of a tubular organ and connecting two points on the circumference.
Outer Diameter | C96684 | Kilometer | km | The greatest possible length of a straight line passing through the center of a circular or spheroid object that connects two points on the circumference.
Pore Size | C112332 | Meter | m | A quantitative or qualitative measurement of the physical dimensions of the pores in a material.
Width | C25345 | Micrometer | um | The extent or measurement of something from side to side.
|  | Millimeter | mm
|  | Nanometer | nm
|  | Picometer | pm
|  | Yard | [yd_i]
Pressure Units of Measure
Size Types | Size Type C-Code | Description | Code
Pressure | C25195 | Hertz | Hz | The force applied to a unit area of surface.
|  | KiloPascal | kPa
|  | millibar | mbar
|  | Pound per Square Inch | [psi]
Volume Units of Measure
Size Types |  | Description | Code
Total Volume | C25335 | Centiliter | cL | The total amount of three dimensional space occupied by an object or the capacity of a space or container.
|  | Cubic Inch | [cin_i]
|  | Cup | [cup_us]
|  | Deciliter | dL
|  | Femtoliter | fL
|  | Fluid Ounce | [foz_us]
|  | Gallon | [gal_us]
|  | Kiloliter | kL
|  | Liter | L
|  | Microliter | uL
|  | Milliliter | mL
|  | Nanoliter | nL
|  | Picoliter | pL
|  | Pint | [pt_us]
|  | Quart | [qt_us]
Weight Units of Measure
Size Types |  | Description | Code
Weight | C25208 | Gram | g | The vertical force exerted by a mass as a result of gravity.
|  | Kilogram | kg
|  | Metric Ton | t
|  | Microgram | ug
|  | Milligram | mg
|  | Pound | [lb_av]
|  | Ton | [ston_av]
Device Size Text | C106041 |  |  | Additional undefined device size not represented in the GUDID clinically relevant size list.
Storage and Handling Units of Measure
Types |  | Description | Code
Storage Environment Atmospheric Pressure | C101705 | KiloPascal | kPa | Storage requirements that are indicated for the device; atmospheric pressure under which the product should be stored, e.g., 1000 - 1100 hPa.
|  | millibar | mbar
|  | Inches of Mercury | in[Hg]
|  | Millimeter of Mercury | mm[Hg]
Storage Environment Humidity | C101706 | Percent (%) Relative Humidity | Percent | Storage requirements that are indicated for the device; relative humidity in volume percent under which the product should be stored, e.g., 0 - 20 %.
Storage Environment Temperature | C101707 | Degrees Celsius | Cel | Storage requirements that are indicated for the device; environmental temperature range under which the device should be stored.
|  | Degrees Fahrenheit | [degf]
|  | Degrees Kelvin | K
Handling Environment Atmospheric Pressure | C101681 | KiloPascal | kPa | Handling requirements that are indicated for the device; atmospheric pressure under which the product should be operated, e.g., 1000 - 1100 hPa.
|  | millibar | mbar
|  | Inches of Mercury | in[Hg]
|  | Millimeter of Mercury | mm[Hg]
Handling Environment Humidity | C101682 | Percent (%) Relative Humidity | Percent | Handling requirements that are indicated for the device; relative humidity in volume percent under which the product should be operated, e.g., 0 - 20 %.
Handling Environment Temperature | C101683 | Degrees Celsius | Cel | Handling requirements that are indicated for the device; environmental temperature range under which the device may be operated.
|  | Degrees Fahrenheit | [degf]
|  | Degrees Kelvin | K
Special Storage Conditions | C101704 |  |  | Any special storage requirements for this device.
Sterilization Method
Types
Chlorine Dioxide (ClO2) Sterilization | C101691 |  |  | A sterilization process that uses chlorine dioxide (ClO2) gas at ambient temperatures and low concentrations to kill microorganisms.
Dry Heat Sterilization | C101712 |  |  | A sterilization process that uses high temperature and dry air to kill microorganisms.
Ethylene Oxide Sterilization | C101713 |  |  | A sterilization process that uses ethylene oxide (EtO) gas to kill microorganisms.
High Intensity Light or Pulse Light Sterilization | C101693 |  |  | A sterilization process that uses intense and short duration pulses of wavelengths of light from the ultraviolet to the near infrared regions to kill microorganisms.
High-level Disinfectant Sterilization | C107419 |  |  | A germicide that inactivates all microbial pathogens, except small numbers of bacterial endospores.
Hydrogen Peroxide (H2O2) Sterilization | C101689 |  |  | A sterilization process that uses radicals that form from hydrogen peroxide (H2O2) to kill microorganisms.
Liquid Chemical Sterilization | C107418 |  |  | A sterilization process that uses liquid chemical as its sterilant to kill microorganisms.
Microwave Radiation Sterilization | C101694 |  |  | A sterilization process that uses nonionizing radiation at frequencies between 300MHz and 300GHz in the microwave region to kill microorganisms.
Moist Heat or Steam Sterilization | C101714 |  |  | A sterilization process that uses moist air at high temperatures to kill microorganisms.
Nitrogen Dioxide Sterilization | C107416 |  |  | A sterilization process that uses nitrogen dioxide radicals that are formed from dinitrogen tetroxide dimers to kill microorganisms.
Ozone (O3) Sterilization | C101690 |  |  | A sterilization process that uses radicals that form from ozone(O3) gas to kill microorganisms.
Peracetic Acid Sterilization | C101697 |  |  | A sterilization process that uses the vaporized form of a chemical (peracetic acid) to kill microorganisms.
Radiation Sterilization | C101715 |  |  | A sterilization process that uses radiation such as electron beams, X-rays, gamma rays, or subatomic particles to kill microorganisms.
Sound Waves Sterilization | C101695 |  |  | A sterilization process that uses high-frequency sound waves to kill microorganisms.
Sterilization Method | C84382 |  |  | Validated process used to render product free from viable microorganisms. See ISO/TS 11139.D7
Supercritical Carbon Dioxide Sterilization | C107417 |  |  | A sterilization process that uses supercritical carbon dioxide in combination with a small percentage of additive to kill microorganisms.
Ultraviolet Light Sterilization | C101696 |  |  | A sterilization process that uses short wavelength radiation to kill microorganisms.
MR Safety
Types | C-Code
Labeling does not contain MRI Safety Information | C113844 |  |  | Information on the safety of using a device in a magnetic resonance (MR) environment has not been provided in the labeling by the labeler. This information is not required by FDA regulation or the UDI Rule.
MR Conditional | C106046 |  |  | A device with demonstrated safety in the MR environment within defined conditions. At a minimum, address the conditions of the static magnetic field, the switched gradient magnetic field and the radiofrequency fields. Additional conditions, including specific configurations of the item, may be required. See ASTM F2503. Reprinted, with permission, from F2503-08 Standard Practice for Marking Medical Devices and Other Items for Safety in the Magnetic Resonance Environment, copyright ASTM International, 100 Barr Harbor Drive, West Conshohocken, PA 19428.
MR Safe | C106045 |  |  | The device poses no known hazards resulting from exposure to any MR environment. MR Safe items are composed of materials that are electrically nonconductive, nonmetallic, and nonmagnetic. See ASTM F2503. Reprinted, with permission, from F2503-08 Standard Practice for Marking Medical Devices and Other Items for Safety in the Magnetic Resonance Environment, copyright ASTM International, 100 Barr Harbor Drive, West Conshohocken, PA 19428.
MR Unsafe | C106047 |  |  | The device which poses unacceptable risks to the patient, medical staff or other persons within the MR environment. See ASTM F2503. Reprinted, with permission, from F2503-08 Standard Practice for Marking Medical Devices and Other Items for Safety in the Magnetic Resonance Environment, copyright ASTM International, 100 Barr Harbor Drive, West Conshohocken, PA 19428.
MRI Safety | C106044 |  |  | Magnetic Resonance Imaging (MRI) Safety indicates that sufficient testing has been conducted to characterize the behavior of the device in the magnetic resonance (MR) environment. See ASTM F2503.
Footnotes
1 | This list of values (LOV) is current as of the release of this document. As new LOVs are identified, the production environment will be updated to accept additional code values.
2 | C-Code is the NCIt concept code attached to the concept. NCIt Codes are unique strings that begin with a C and are followed by a series of digits. See Helpful Links sheet for link to download the latest version of C-codes (NCIt Concept codes).
3 | See Helpful Links sheet for latest for UCUM (Unified Code for Units of Measure) Units of Measure Codes

## Sample Label
Sample Label
Fictitious Medical Device Label
(“Not all fields in GUDID are required to be on the label. This example is for illustrative purpose only”)
Figure 1: Fictitious Medical Device label

## Helpful Links
Helpful Links
FDA UDI main page | https://www.fda.gov/udi
FDA UDI Help Desk | https://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/UniqueDeviceIdentification/ucm368904.htm
FDA UDI Resources | https://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/UniqueDeviceIdentification/ChangesbetweenUDIProposedandFinalRules/default.htm
AccessGUDID | https://accessgudid.nlm.nih.gov/ | Public portal for the GUDID data submitted to the FDA by the labelers
OpenFDA | https://open.fda.gov | FDA website that allows collaboration and linkage of UDI data to other CDRH data sets.
FDA UDI Optimizing Data Quality | https://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/UniqueDeviceIdentification/ucm596243.htm | Current priorities on data quality and links to early adopters
FDA UDI GUDID Resources | https://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/UniqueDeviceIdentification/GlobalUDIDatabaseGUDID/default.htm | GUDID resources page with additional links to other GUDID resources
ONC UDI implantables list 2015 edition certification - 45 CFF 170.315(a)(14) | https://www.healthit.gov/test-method/implantable-device-list | Relevant to Footnote 6 in GUDID Data Elements Sheet
CDRH GUDID C-Codes from NCIt | https://evs.nci.nih.gov/ftp1/FDA/CDRH/GUDID/About.html | Relevant to C-Codes included in the List of Values sheet
UCUM Units of Measure Codes | http://unitsofmeasure.org/ucum.html | Relevant to units of measures included in the List of Values sheet
GMDN Agency | https://www.gmdnagency.org/ | Relevant to GMDN code, terms and description data elements