## USFDA
Code (TOC Level) | Level | Scope | DisplayName | ORIGINAL PMA | PTS | 180-day N | 180-day O | Real-Time | 30-day notice/135-day replacment | Condition (if applicable) | Traditional 510(k) | Abbreviated 510(k) | Special 510(k) | Condition (if applicable)
CHAPTER 1 – REGIONAL ADMINISTRATIVE
CH1.01 | 1 | IMDRF | Cover Letter | CR | CR | CR | CR | CR | CR | Required if no FDA Form 3514 is provided. Preference is to have both. | CR | CR | CR | Required if no FDA Form 3514 is provided. Preference is to have both.
CH1.02 | 1 | IMDRF | Submission ToC | R | R | O | O | R | R |  | R | R | R
CH1.03 | 1 | IMDRF | List of Terms/ Acronyms | OR | OR | OR | OR | OR | OR |  | OR | OR | OR
CH1.04 | 1 | IMDRF, RF | Application Form/Administrative Information | CR | CR | CR | CR | CR | CR | Required if no Cover Letter is provided. Preference is to have both. | CR | CR | CR | Required if no Cover Letter is provided. Preference is to have both. Note use of form 3414 is voluntary
CH1.05 | 1 | IMDRF RF (Anvisa, EU, USFDA, TGA) | Listing of Device Models/Variants | R | R | R | R | R | R |  | R | R | R
CH1.06 | 1 | Regional (HC, Anvisa, EU, TGA) | Quality Management System, Full Quality System or Product Certification Certificate | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH1.07 | 1 | Regional (Anvisa) | Free Sale Certificate | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH1.08 | 1 | Regional (USFDA, Anvisa, EU) | User Fees | R | R | R | NR | R | R |  | R | R | R
CH1.09 | 1 | IMDRF, RF | Pre-Submission Correspondence and Previous Regulator Interactions | R | R | CR | CR | R | CR | When presubmission interactions have accured, describe those interactions. | R | R | R
CH1.10 | 1 | Regional (USFDA) | Acceptance for Review Checklist | O | O | NR | NR | NR | NR |  | O | O | O
CH1.11 | 1 | Regional (USFDA, Anvisa, EU, TGA) | Statements/Certifications/Declarations of Conformity | - | [6 empty] | -
CH1.11.1 | 2 | Regional (USFDA) | Performance and Voluntary Standard | R | R | O | O | O | O |  | O | CR | O | If the Abbreviated Declares Conformity with Standards
CH1.11.2 | 2 | Regional (USFDA) | Environmental Assessment | R | R | NR | NR | NR | NR |  | NR | NR | NR
CH1.11.3 | 2 | Regional (USFDA) | ClinicalTrials.gov | R | R | CR | NR | NR | NR | Required when submitting investigational clinical data. | CR | CR | CR | Required when submitting clinical data
CH1.11.4 | 2 | Regional (USFDA) | Indications for Use Statement with Rx and OTC designation Enclosure | R | R | R | NR | NR | NR |  | R | R | R
CH1.11.5 | 2 | Regional (USFDA, Anvisa, TGA) | Truthful and Accurate Statement | NR | NR | NR | NR | NR | NR |  | R | R | R
CH1.11.6 | 2 | IMDRF (EU, TGA) | Declaration of Conformity | CR | CR | CR | CR | CR | CR |  | CR | CR | R | See CH 3.4.2 for Dec of Conformity that is required
CH1.12 | 1 | IMDRF | Letters of Reference for Master Files | CR | CR | CR | CR | CR | CR | required when referencing a master file | CR | CR | CR | required when referencing a master file
CH1.13 | 1 | Regional (Anvisa) | Letter of Authorization | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH1.14 | 1 | IMDRF | Other Regional Administrative Information | CR | CR | CR | CR | CR | CR | When information is requested by the regulator (through guidance documents or other communication) but does not belong in any of the other submission context headings. | CR | CR | CR | When information is requested by the regulator (through guidance documents or other communication) but does not belong in any of the other submission context headings.
CHAPTER 2 – SUBMISSION CONTEXT
CH2.1 | 1 | IMDRF | Chapter ToC | OR | OR | OR | OR | OR | OR |  | OR | OR | OR
CH2.2 | 1 | IMDRF | General Summary of Submission | O | O | OR | OR | OR | OR |  | O | O | O
CH2.3 | 1 | Regional (USFDA) | Summary and Certifications for Premarket Submissions | R | R | O | O | O | O |  | R | R | R
CH2.4 | 1 | IMDRF | Device Description | - | [6 empty] | - | - | -
CH2.4.1 | 2 | IMDRF, RF | Comprehensive Device Description & Principle of Operation | R | R | CR | CR | CR | CR | Include device description in supplements if device is changing and the description should focus on aspect of device that is changing | R | R | R
CH2.4.2 | 2 | IMDRF (HC, JP) | Material Specifications | - | [6 empty] | - | - | -
CH2.4.3 | 2 | TGA, USFDA | Description of Device Packaging | R | R | R | CR | CR | CR | Required if supplement includes a change to packaging or sterilization | R | R | CR | Required per 510(k) RTA guidance document if Sterilziation Validation information was provided.
CH2.4.4 | 2 | IMDRF (ANVISA, EU, HC, TGA, USFDA) | History of Development | OR | OR | OR | O | OR | O |  | O | O | O
CH2.4.5 | 2 | IMDRF, RF | Reference and Comparison to Similar and/or Previous Generations of the Device | O | CR | CR | NR | CR | NR | Required for device design changes | O | O | O
CH2.4.6 | 2 | Regional (USFDA) | Substantial Equivalence Discussion | NR | NR | NR | NR | NR | NR |  | R | R | R
CH2.5 | 1 | IMDRF, RF | Indications for Use and/or Intended Use and Contraindications | - | [6 empty] | - | - | -
CH2.5.1 | 2 | IMDRF, RF | Intended Use / Intended Purpose / Intended User / Indications for Use | R | R | CR | NR | NR | NR | Required if indications for use are changing | R | R | R
CH2.5.2 | 2 | IMDRF, RF | Intended Environment/Setting for use | CR | CR | CR | NR | CR | NR | If the environment is important in the use of the device or this aspect is changed | CR | CR | CR | If the environment is important in the use of the device.
CH2.5.3 | 2 | Regional (USFDA) | Pediatric Use | R | R | R | R | R | NR | Required to submit info or provide a statement in accordance with Section 515A of the FD&C Act | R | R | R
CH2.5.4 | 2 | IMDRF, RF | Contraindications For Use | CR | CR | CR | NR |  |  | When the medical device has contraindications or this section of the labeling is being updated in the supplement | CR | CR | CR | when the medical device has contraindications
CH2.6 | 1 | IMDRF | Global Market History | - | [6 empty] | - | - | -
CH2.6.1 | 2 | IMDRF | Global Market History | R | R | NR | NR | NR | NR |  | O | O | O
CH2.6.2 | 1 | IMDRF | Global Incident Reports and Recalls | R | R | O | NR | NR | NR |  | CR | CR | CR | Required for 510(k) submission utilizing Blue Book memorandum #K95-1
CH2.6.3 | 2 | IMDRF (HC, JP, EU, TGA) | Sales, Incident and Recall Rates | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH2.6.4 | 2 | Regional (TGA) | Evaluation/Inspection Reports | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH2.7 | 1 | IMDRF | Other Submission Context Information | CR | CR | CR | CR | CR | CR | Identify if changes are related to field events, recalls, complaints, etc. | CR | CR | CR | If additional information is needed
CHAPTER 3 – NON-CLINICAL EVIDENCE
CH3.1 | 1 | IMDRF | Chapter ToC | OR | OR | OR | NR | NR | NR |  | OR | OR | OR
CH3.2 | 1 | IMDRF | Risk Management | R | R | O | NR | O | NR |  | O | O | R
CH3.3 | 1 | IMDRF (EU, TGA ) | Essential Principles (EP) Checklist | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH3.4 | 1 | IMDRF, RF | Standards | - | [6 empty] | - | - | -
CH3.4.1 | 2 | IMDRF, RF | List of Standards | CR | CR | CR | O | O | NR | If standards are used in support of approval | CR | CR | CR | If standards are used in support of Substantial Equivalence
CH3.4.2 | 2 | Regional (HC, Anvisa, USFDA) | Declaration and/or Certification of Conformity | CR | CR | CR | O | O | NR | Required only when an applicant choose to declare conformity | O | CR | R | If the Abbreviated declares conformity with standards (i.e., the sponsor is not only including a guidance Summary Report)
CH3.5 | 1 | IMDRF | Analytical Performance | - |  |  |  | R |  |  | - | - | -
CH3.5.01 | 2 | IMDRF | Stability of Sample(s) | CR | CR | CR | NR | CR | NR | Required only when submission is for a new assay and/or new type of specimens | CR | CR | CR | If this information is different in the Special in comparison to the originally cleared device.
CH3.5.01.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | CRI | CRI | CRI | NR | CR | NR | When any testing/studies is conducted in support of the submission under the parent heading subject, this is required for each study/test. | CRI | CRI | CR | When any testing/studies is conducted in support of the submission under the parent heading subject, this is required for each study/test.
CH3.5.01.1.1 | 4 | IMDRF | Summary | CR | CR | CR | NR | CR | NR | When any testing/studies is conducted in support of the submission under the parent heading subject, this is required for each study/test. | CR | CR | CR | When any testing/studies is conducted in support of the submission under the parent heading subject, this is required for each study/test.
CH3.5.01.1.2 | 4 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR | When any testing/studies is conducted in support of the submission under the parent heading subject, this is required for each study/test. These reports may be requested if not submitted following the review of the test/study summary. | CR | CR | CR | When any testing/studies is conducted in support of the submission under the parent heading subject, this is required for each study/test. These reports may be requested if not submitted following the review of the test/study summary.
CH3.5.01.1.3 | 4 | Regional (USFDA) | Statistical Data | CR | CR | CR | NR | CR | NR | If studies require it | CR | CR | CR
CH3.5.02 | 2 | IMDRF | Validation of Samples | R | R | R | NR | CR | NR | Required if submission is for a new design/material | R | R | CR | If this information is different in the Special in comparison to the originally cleared device.
CH3.5.02.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | CRI | CRI | CRI | NR | CR | NR |  | CRI | CRI | CR
CH3.5.02.1.1 | 4 | IMDRF | Summary | CR | CR | CR | NR | CR | NR | When any testing/studies is conducted in support of the submission under the parent heading subject, this is required for each study/test. These reports may be requested if not submitted following the review of the test/study summary. | CR | CR | CR
CH3.5.02.1.2 | 4 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR |  | CR | CR | NR
CH3.5.02.1.3 | 4 | IMDRF | Statistical Data | CR | CR | CR | NR | CR | NR |  | CR | CR | NR
CH3.5.03 | 2 | IMDRF | Metrological traceability of calibrator and control material values | R | R | R | NR | CR | NR |  | R | CR | CR | Not needed if this information is already established in a marketed product.
CH3.5.03.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | R | CR | CR
CH3.5.03.1.1 | 4 | IMDRF | Summary | R | R | R | NR | CR | NR |  | CR | CR | CR
CH3.5.03.1.2 | 4 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | CR | CR | CR
CH3.5.03.1.3 | 5 | IMDRF | Statistical Data | CR | CR | CR | CR | CR | CR |  | CR | CR | CR
CH3.5.04 | 2 | IMDRF | Accuracy of Measurement |  |  |  |  | CR | NR
CH3.5.04.1 | 3 | IMDRF | Trueness | CR | CR | CR | NR | CR | NR | If reference method, reference standard, traceable methods exist | CR | CR | NR | Method comparison studies, recovery, inclusivity CLSI EP09, EP12. Method Comparison study is common
CH3.5.04.1.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | R | CR | CR
CH3.5.04.1.1.1 | 4 | IMDRF | Summary | R | R | R | NR | CR | NR |  | R | CR | CR
CH3.5.04.1.1.2 | 4 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | R | OR | NR
CH3.5.04.1.1.3 | 5 | Regional (USFDA) | Statistical Data | R | R | R | NR | CR | NR |  | R | CR | NR | The applicant is advised to contact the specific review division for further guidance on the specific data format that is preferred
CH3.5.04.2 | 3 | IMDRF | Precision (Repeatability and Reproducibility) | R | R | R | NR | CR | NR |  | R | R | NR | EP05
CH3.5.04.2.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | R | R | NR
CH3.5.04.2.1.1 | 5 | IMDRF | Summary | R | R | R | NR | CR | NR |  | R | R | R
CH3.5.04.2.1.2 | 5 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | OR | OR | NR
CH3.5.04.2.1.3 | 5 | Regional (USFDA) | Statistical Report | R | R | R | NR | CR | NR |  | OR | OR | NR
CH3.5.05 | 3 | IMDRF | Analytical Sensitivity | R | R | R | NR | CR | NR |  | R | R | CR | LoB, LoD, LoQ studies, EP17, EP12
CH3.5.05.1 | 4 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | CR | CR | CR
CH3.5.05.1.1 | 5 | IMDRF | Summary | R | R | R | NR | CR | NR |  | CR | CR | CR
CH3.5.05.1.2 | 5 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | CR | CR | NR
CH3.5.05.1.3 | 5 | Regional (USFDA) | Statistical Data | R | R | R | NR | CR | NR |  | CR | CR | NR
CH3.5.06 | 2 | IMDRF | Analytic Specificity | R | R | R | NR | CR | NR |  | R | R | CR | Interferences and cross-reactivity studies, EP07. Analyte/measurand levels should challenge medical decision points for interference testing. Analytes/substances tested for crossreactivity and/or interference should be tested at the highest level that can be expected in a clinical specimen
CH3.5.06.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | R | R | CR
CH3.5.06.1.1 | 4 | IMDRF | Summary | R | R | R | NR | CR | NR |  | R | R | CR
CH3.5.06.1.2 | 4 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | R | R | NR
CH3.5.06.1.3 | 4 | Regional (USFDA) | Statistical Data | R | R | R | NR | CR | NR |  | R | R | NR
CH3.5.07 | 2 | IMDRF | Hook Effect see Measuring Interval of the Assay (CH3.5.08) | CR | CR | CR | NR | CR | NR |  | CR | CR | CR | Needed when device type is known to demonstrate effect or first of a kind. If this information is different in the Special in comparison to the originally cleared device.
CH3.5.07.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.5.07.1.1 | 4 | IMDRF | Summary | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.5.07.1.2 | 4 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR |  | CR | CR | NR
Ch3.5.07.1.3 | 4 | Regional (USFDA) | Statistical Data | CR | CR | CR | NR | CR | NR |  | Cr | CR | NR
CH3.5.08 | 2 | IMDRF | Measuring Interval of the Assay | R | R | R | NR | CR | NR |  | CR | CR | CR | Studies as linearity, hook effect, CLSI EP06.
CH3.5.08.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | R | R | CR
CH3.5.08.1.1 | 4 | IMDRF | Summary | R | R | R | NR | CR | NR |  | R | R | R
CH3.5.08.1.2 | 4 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | R | R | NR
CH3.5.08.1.3 | 5 | Regional (USFDA) | Statistical Data | R | R | R | NR | CR | NR |  | R | R | NR
CH3.5.09 | 2 | IMDRF | Validation of the Assay Cut-off | R | R | R | NR | CR | NR |  | R | R | CR | If this information is different in the Special in comparison to the originally cleared device.
CH3.5.09.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | R | R | CR
CH3.5.09.1.1 | 4 | IMDRF | Summary | R | R | R | NR | CR | NR |  | R | R | R
CH3.5.09.1.2 | 4 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | R | R | NR
CH3.5.09.1.3 | 5 | Regional (USFDA) | Statistical Data | R | R | R | NR | CR | NR |  | R | R | NR
CH3.5.10 | 2 | IMDRF | Validation of the Assay Procedure | O | O | O | NR | CR | NR |  | O | O | O
CH3.5.10.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | R | R | R | NR | CR | NR |  | R | R | CR
CH3.5.10.1.1 | 4 | IMDRF | Summary | R | R | R | NR | CR | NR |  | R | R | CR
CH3.5.10.1.2 | 4 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | OR | OR | NR
CH3.5.10.1.3 | 4 | Regional (USFDA) | Statistical Data |  |  |  | NR | CR | NR |  | CR | CR | CR
CH3.6 | 1 | IMDRF | Other Studies |  |  |  | NR
CH3.6.1 | 2 | IMDRF | Electrical Safety, Mechanical and Environmental Protection, and Electromagnetic Compatibility | R | R | R | NR | CR | NR |  | R | R | CR | This information is different in the Special in comparison to the originally cleared device.
CH3.6.1.1 | 3 | IMDRF | [Study description, study identifier, date of initiation, date of completion] | CRI | CRI | CRI | NR | CR | NR |  | CRI | CRI | CRI
CH3.6.1.1.1 | 4 | IMDRF | Summary | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.1.1.2 | 4 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.1.1.3 | 4 | Regional (USFDA) | Statistical Data | NR | NR | NR | NR | CR | NR |  | NR | NR | NR
CH3.6.2 | 2 | IMDRF | Software/Firmware | CR | CR | CR | NR | CR | NR |  | CR | CR | CR | If the device contains software/firmware. This information is different in the Special in comparison to the originally cleared device.
CH3.6.2.01 | 3 | IMDRF | Software/Firmware Description | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.02 | 3 | IMDRF | Hazard Analysis | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.03 | 3 | IMDRF | Software Requirement Specification | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.04 | 3 | IMDRF (EU, HC, JP, USFDA) | Architecture Design Chart | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.05 | 3 | IMDRF (EU, HC, JP, USFDA) | Software Design Specification | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.06 | 3 | IMDRF | Traceability Analysis | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.07 | 2 | IMDRF | Software Life Cycle Process Description | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.08 | 3 | IMDRF | Software Verification and Validation | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.08.1 | 4 | IMDRF | [Study description, study identifier, date of initiation] | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.2.08.1.1 | 5 | IMDRF | Summary | CR | CR | CR | NR | CR | NR |  | CR | CR | NR
CH3.6.2.08.1.2 | 5 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR |  | CR | CR | NR
CH3.6.25.08.1.3 | 5 | Regional (USFDA) | Statistical Data | NR | NR | NR | NR | CR | NR |  | NR | NR | NR
CH3.6.2.09 | 3 | IMDRF | Revision Level History |  |  |  | NR | CR | NR |  | R | R | CR
CH3.6.2.10 | 3 | IMDRF | Unresolved Anomalies (Bugs or Defects) |  |  |  | NR | CR | NR |  | CR | CR | CR
CH3.6.3 | 2 | IMDRF | Cleaning and Disinfection Validation | CR | CR | CR | NR | CR | NR |  | CR | CR | CR | If the specific device requires this for safe use. If this information is different in the Special in comparison to the originally cleared device.
CH3.6.3.1 | 3 | IMDRF | [Study description, study identifier, date of initiation] | CRI | CR | CR | NR | CR | NR |  | CRI | CRI | CRI
CH3.6.3.1.1 | 4 | IMDRF | Summary | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.3.1.2 | 4 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.3.1.3 | 5 | Regional (USFDA) | Statistical Data | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.4 | 2 | IMDRF | Usability/Human Factors | CR | CR | CR | NR | CR | NR |  | R | R | CR | May be needed if OTC use. If this information is different in the Special in comparison to the originally cleared device.
CH3.6.4.1 | 3 | IMDRF | [Study description, study identifier, date of initiation] | CRI | CR | CR | NR | CR | NR |  | CRI | CRI | CRI
CH3.6.1.1 | 4 | IMDRF | Summary | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.4.1.2 | 4 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.4.1.3 | 4 | Regional (USFDA) | Statistical Data |  |  |  | NR | CR | NR |  | CR | CR | CR
CH3.6.5 | 2 | IMDRF | Stability of the IVD | R | R | R | NR | CR | NR |  | R | R | CR | If this information is different in the Special in comparison to the originally cleared device. Information needs to be generated by sponsor, but does not need to be submitted in the 510k.
CH3.6.5.1.1 | 3 | IMDRF | [Study description, study identifier, date of initiation] | R | R | R | NR | CR | NR |  | R | CRI | CRI
CH3.6.5.1.1.1 | 4 | IMDRF | Summary | R | R | R | NR | CR | NR |  | R | CR | CR
CH3.6.5.1.1.2 | 4 | IMDRF | Full Report | R | R | R | NR | CR | NR |  | R | CR | CR
CH3.6.5.1.1.3 | 4 | Regional (USFDA) | Statistical Data |  |  |  | NR | CR | NR |  | R | CR | CR
CH3.6.5.2 | 3 |  | In Use Stability | R | CR | CR | NR | CR | NR |  | R | CR | CR | EP 25 Needed when reagents are not stable at room conditions
CH3.6.5.2.1 | 4 |  | [Study description, study identifier, date of initiation] | R | CR |  | NR | CR | NR |  | CR | CR | CR | If this information is different in the Special in comparison to the originally cleared device.
CH3.6.5.2.1.1 | 5 | IMDRF | Summary | R | CR | CR | NR | CR | NR |  | CR | CR | NR
CH3.6.5.2.1.2 | 5 | IMDRF | Full Report | R | R | CR | NR | CR | NR |  | CR | CR | NR
CH3.6.5.2.1.3 | 5 | Regional (USFDA) | Statistical Data | R | CR | NR | NR | CR | NR |  | NR | NR | NR
CH3.6.5.3 | 3 | IMDRF | Shipping Stability | R | CR | CR | NR | CR | NR |  | CR | CR | CR | Needed when reagents/components are not stable at ambient conditions
CH3.6.5.3.1 | 4 | IMDRF | [Study description, study identifier, date of initiation] | R | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.5.3.1.1 | 5 | IMDRF | Summary | R | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.6.5.3.1.2 | 5 | IMDRF | Full Report | R | CR | CR | NR | CR | NR |  | CR |  |  | If this information is different in the Special in comparison to the originally cleared device.
CH3.6.5.3.1.3 | 5 | IMDRF | Statistical Data | R | CR | CR | NR | CR | NR |  | CR | CR | NR
CH3.7 | 1 | IMDRF | Analytical Performance and other Evidence Bibliography | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
3.8 | 1 | IMDRF | Other Evidence |  | CR | CR | NR | CR | NR
CH3.8.1 | 2 | IMDRF | [Study description, study identifier, date of initiation] | CRI | CR | CR | NR | CR | NR |  | CRI | CRI | CRI
CH3.8.1.1 | 3 | IMDRF | Summary | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.8.1.2 | 3 | IMDRF | Full Report | CR | CR | CR | NR | CR | NR |  | CR | CR | CR
CH3.8.1.3 | 3 | Regional (USFDA) | Statistical Data |  |  |  | NR | CR | NR |  | CR | CR | CR
CHAPTER 4 – CLINICAL EVIDENCE
CH4.1 | 1 | IMDRF | Chapter ToC | OR | OR | OR | OR | NR | NR |  | OR | OR | NR
CH4.2 | 1 | IMDRF | Overall Clinical Evidence Summary | R | R | CR | NR | NR | NR | Required when clinical data is submitted | CR | CR | NR | required when clinical data is supplied
CH4.2.1 | 2 | IMDRF | Expected Values/Rference Intervals | R | CR | CR | CR | NR | NR |  | R | R | R | EP28
CH4.2.2 | 2 | Regional (TGA, EU) | Clinical Evidence Evaluation Report | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH4.2.3 | 2 | IMDRF | Device Specific Clinical Studies | - | [6 empty] | - | - | NR
CH4.2.3.1 | 3 | IMDRF | [Study description, protocol#, date of initiation, date of completion] | CRI | CRI | CRI | CRI | NR | NR |  | CRI | CRI | NR
CH4.2.3.1.1 | 4 | IMDRF | Clinical Study Synopsis | R | CR | CR | CR | NR | NR | Required when clinical data is submitted | CR | CR | NR | required when clinical data is submitted
CH4.2.3.1.2 | 4 | IMDRF | Clinical Study Report | R | CR | CR | CR | NR | NR | Required when clinical data is submitted | CR | CR | NR | required when clinical data is supplied
CH4.2.3.1.3 | 4 | Regional (USFDA) | Clinical Study Data | R | CR | CR | CR | NR | NR | Required when clinical data is submitted | CR | CR | NR | required when noted in regulations or guidance documents
CH4.2.3 | 2 | IMDRF | Clinical Literature Review and Other Reasonable Known Information | R | R | CR | CR | NR | NR | Required when clinical data is submitted | CR | CR | NR | required when clinical literature is submitted
CH4.3 | 1 | Regional ( USFDA) | IRB Approved Informed Consent Forms | R | O | NR | NR | NR | NR |  | NR | NR | NR
CH4.4 | 1 | Regional (USFDA) | Investigators Sites and IRB contact information | R | CR | NR | NR | NR | NR | Required when clinical data is submitted | NR | NR | NR
CH4.5 | 2 | IMDRF | Other Clinical Evidence | - |  |  |  | NR | NR |  | - | - | NR
CH4.5.1 | 3 | IMDRF | [Study description, study identifier, date of initiation] | CRI | CRI | CRI | CRI | NR | NR |  | CRI | CRI | NR
CH4.5.1.1 | 4 | IMDRF | Summary | CR | CR | O | O | NR | NR | Required when clinical data is submitted | CR | CR | NR
CH4.5.1.2 | 4 | IMDRF | Full Report | CR | CR | O | O | NR | NR | When any testing/studies are conducted in support of the submission under the parent heading subject, this is required for each study/test. These reports may be requested if not submitted following the review of the test/study summary. | CR | CR | NR
CH4.5.1.3 | 4 | Regional (USFDA) | Statistical Data | CR | CR | O | O | NR | NR |  | CR | CR | NR
CH4.6 | 1 | Regional (USFDA) | Location of clinical study records | CR | CR | O | O | NR | NR | If studies require it | NR | NR | NR
CHAPTER 5 – LABELLING AND PROMOTIONAL MATERIAL
CH5.01 | 1 | IMDRF | Chapter ToC | O | O | O | O | O | NR |  | O | O | O
CH5.02 | 1 | IMDRF, RF | Product/Package Labels | R | R | CR | CR | CR | NR | If submission includes a change to the label | R | R | R
CH5.03 | 1 | IMDRF, RF | Package Insert/Instructions for Use | R | R | CR | CR | CR | NR | If submission includes a change to the insert/instructions for use | R | R | R
CH5.04 | 1 | IMDRF (EU, Anvisa) | e-labeling | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH5.05 | 1 | Regional (USFDA) | Patient Labelling | R | R | CR | CR | CR | NR | If submission includes a change to the physician labeling | CR | CR | CR | When the device type utilizes this type of labeling
CH5.06 | 1 | IMDRF | Technical/Operators Manual | R | R | CR | CR | CR | NR | If submission includes a change to the technical/operators manual | CR | CR | CR | When the device type utilizes this type of labeling
CH5.07 | 1 | Regional (HC) | Product Brochures | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH5.8 | 1 | IMDRF | Other Labelling and Promotional Material | CR | CR | CR | CR | CR | CR | If there is other labeling associted with the device | CR | CR | CR
CHAPTER 6A – QUALITY MANAGEMENT SYSTEM PROCEDURES
CH6A.1 | 1 | Regional (USFDA) | Cover Letter | R | O | O | O | O | O |  | NR | NR | NR
CH6A.2 | 1 | IMDRF | Chapter ToC | R | O | O | O | O | O |  | NR | NR | NR
CH6A.3 | 1 | IMDRF | Administrative | - | [6 empty] | - | - | -
CH6A.3.1 | 2 | IMDRF | Product Descriptive Information | R | CR | CR | CR | CR | CR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.3.2 | 2 | IMDRF (USFDA, Anvisa, TGA), RF | General Manufacturing Information | R | CR | CR | CR | CR | CR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.3.3 | 2 | IMDRF, RF | Required Forms | CR | CR | CR | CR | CR | CR | Provide only the information specific to the manfacturing change | CR | CR | CR
CH6A.4 | 1 | IMDRF | Quality management system procedures | R | CR | CR | CR | CR | O | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.5 | 1 | IMDRF | Management responsibilities procedures | R | CR | NR | CR | NR | NR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.6 | 1 | IMDRF | Resource management procedures | R | CR | CR | CR | NR | NR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.7 | 1 | IMDRF | Product realization procedures | R | CR | CR | CR | CR | CR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.7.1 | 2 | IMDRF | Design and development procedures | R | CR | CR | CR | NR | NR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.7.2 | 2 | IMDRF | Purchasing procedures | R | CR | CR | CR | CR | CR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.7.3 | 2 | IMDRF | Production and service controls procedures | R | CR | CR | CR | CR | O | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.7.4 | 2 | IMDRF | Control of monitoring and measuring devices procedures | R | CR | CR | CR | CR | O | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.8 | 1 | IMDRF | QMS measurement, analysis and improvement procedures | R | CR | CR | CR | NR | NR | Provide only the information specific to the manfacturing change | NR | NR | NR
CH6A.9 | 1 | IMDRF | Other Quality System Procedures Information | CR | CR | CR | CR | CR | CR | Provide only the information specific to the manfacturing change | CR | CR | CR
CHAPTER 6B – QUALITY MANAGEMENT SYSTEM DEVICE SPECIFIC INFORMATION
CH6B.1 | 1 | IMDRF | Chapter ToC | R | O | O | O | O | O |  | NR | NR | NR
CH6B.2 | 1 | IMDRF | Quality management system information | CR | CR | CR | CR | CR | CR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.3 | 1 | IMDRF | Management responsibilities information | CR | CR | CR | CR | CR | CR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.4 | 1 | IMDRF | Resource management information | CR | CR | CR | CR | CR | CR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.5 | 1 | Regional (HC) | Device Specific Quality Plan | NR | NR | NR | NR | NR | NR |  | NR | NR | NR
CH6B.6 | 1 | IMDRF | Product realization information | CR | CR | CR | CR | CR | CR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.6.1 | 2 | Regional (USFDA, Anvisa, TGA) | Design and development information | R | CR | CR | CR | CR | NR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.6.2 | 2 | IMDRF | Purchasing information | CR | CR | CR | CR | CR | CR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.6.3 | 2 | Regional (USFDA, HC, Anvisa, TGA) | Production and service controls information | R | CR | CR | CR | CR | CR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.6.4 | 2 | IMDRF | Control of monitoring and measuring devices information | R | CR | CR | CR | CR | CR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.7 | 1 | IMDRF | QMS measurement, analysis and improvement information | CR | CR | CR | CR | CR | NR | When applicable to the change or called for in the guidance document | NR | NR | NR
CH6B.8 | 1 | IMDRF | Other Device Specific Quality Management System Information | CR | CR | CR | CR | CR | CR | When called for in a guidance document | NR | NR | NR
LEGEND
R = Required CR = Conditionally Required O = Optional
OR = Optional but recommended NR = Not Required - = No Content (requirement should be based on the sub heading classification)
RI = Required for ToC Implementation
CRI = Conditionally Required for ToC Implementation