## Slide 1: Click to add title
Jeffrey R. Miller, MD, MPH
c-jeffmil@pa.gov
CDR United States Public Health Service
CDC Career Epidemiology Field Officer assigned to
Pennsylvania Department of Health
Field Services Branch, Division of State and Local Readiness,
Office of Public Health Preparedness and Response, CDC
Non-Tuberculous Mycobacteria and
Heater-Cooler Units —
Pennsylvania, 2015 – 2016
Presented June 2-3, 2016 at
Circulatory System Devices Panel Meeting,
Center for Devices and Radiological Health
Food and Drug Administration

## Slide 2: Important Notes Before Proceeding
I have no conflicts of interest to declare.
Photos in this presentation are representative of findings from multiple hospitals. Do not assume they were taken at a particular hospital.

## Slide 3: Objectives
Describe findings from index investigation
Present public health response to date, including PADOH guidance to facilities and aggregate findings from hospitals in PA
Share on-going challenges

## Slide 4: Initial Notification
July 2015: “Hospital A” reported a cluster of NTM infections among cardiothoracic surgery patients
Recent reports of an association between heater-cooler units (HCUs) and risk for NTM infections
Hospital A replaced their HCUs and attached devices
Pennsylvania Department of Health (PADOH) and Centers for Disease Control and Prevention (CDC) investigated the cluster of infections

## Slide 5: Investigation Objectives
Determine the extent of the outbreak
Determine associated risk factors and exposures
Recommend control measures

## Slide 6: Investigation Methods
Observed current infection control practices
Reviewed policies and procedures
Conducted epidemiologic study
Analyzed environmental and clinical NTM isolates

## Slide 7: Comparison of Instructions for Use (IFU) for One HCU
Domain | 2010 IFU | 2012 IFU | 2015 IFU
Water type | Sterile or tap water | Filtered water; No DI or RO water | Filtered (0.2 micron) water; No DI or RO water
Water change |  | Should be changed q2weeks | Must be changed weekly
Disinfection | Must disinfect… | Must disinfect… | Must disinfect…
 | Prior to 1st Use | Prior to 1st Use | Prior to 1st Use
 |  | Placed in storage | Placed in storage
 |  | If H2O2 not routinely used | 
 | Regular schedule (e.g., 14 days) | Recommend q3months | Must q14days (in storage or operation)
H2O2 |  | q2weeks with water change | Weekly with water change
 |  | Add 50 ml 3% H2O2 q5days | Use test strips to verify ppm
Tubing and Accessories | Ice container for cooling boost | Include tubes in disinfection; Ice container N/A | Drinking water certified tubing; Specific disinfection steps; Change tubes yearly
Placement | 2.5 ft from walls | 2.5 ft from walls | 2.5 ft from walls; Away from surgical field or Operated out of room
Maintenance | 1000 h or 12 mo | 1000 h or 12 mo | 1000 h or 12 mo
Instructional video includes 56 steps for disinfection, no steps for internal cleaning.

## Slide 8: NTM+ Patients (N = 144) — Hospital A, 2010 – 2015
Characteristic | Patients n (%)
Specimen site | 
Respiratory | 104 (72%)
Invasive (e.g., blood, organ space) | 23 (16%)
Superficial | 17 (12%)
Mycobacterium avium complex (MAC) | 78 (54%)
Had a surgical procedure1 | 48 (33%)
1Between 30 days and 3.5 years prior to NTM+ culture

## Slide 9: NTM+ Cultures by Surgery — Hospital A, 2010 – 2015
Surgery Type | Total | Number of NTM + Patients | Rate per 10,000 surgeries
All surgeries | 99,485 | 48 | 5
Cardiothoracic | 9,899 | 20 | 20
General surgery | 24,572 | 19 | 8
Orthopedic | 15,875 | 8 | 5
The table includes the three most frequently performed surgery categories.
The categories are not mutually exclusive (e.g., one patient could have multiple surgeries).

## Slide 10: Case Control Study at Hospital A
Evaluate risk factors for developing NTM infection after cardiothoracic surgery
Case patients
  Cardiothoracic surgery performed 1/1/2009-7/1/2015
  NTM+ culture 1/1/2010 – 7/1/2015
    Probable: non-respiratory specimen
    Suspect: respiratory specimen
Control patients (3:1 ratio; unmatched design)
  Randomly selected from a list of cardiothoracic surgeries performed between 1/1/2009 – 12/30/2014
Excluded if prior history of NTM infection

## Slide 11: History of Surgery Among Enrolled Patients
 | Probable Cases N = 10 n (%) | Controls N = 48 n (%) | P-value
Surgical History |  |  | 
# of surgeries (any): mean (range) | 2.0 (1 – 5) | 2.3 (1 – 14) | 0.64
# of CT surgeries: mean (range)1 | 1.5 (1 – 3) | 1.8 (1 – 6) | 0.50
# of patients with a non-CT surgery1 | 2 (20) | 11 (23) | 1.0
1Categories are not mutually exclusive

## Slide 12: Demographics of Enrolled Patients
 | Probable Cases N = 10 n (%) | Controls N = 48 n (%) | P-value
Age: mean (range) | 68 (47 – 84) | 64 (20 – 84) | 0.49
Male | 8 (80) | 27 (56) | 0.29
White | 10 (100) | 45 (94) | 1.0
Lung disease | 4 (30) | 8 (17) | 0.19
Diabetes | 1 (10) | 15 (31) | 0.26
Immunocompromised | 3 (30) | 8 (17) | 0.38
HIV | 0 (0) | 0 (0) | —
Transplant | 0 (0) | 0 (0) | —
Chemotherapy | 2 (20) | 9 (19) | 1.0
Steroid treatment | 1 (10) | 2 (4) | 0.44
Other | 3 (30) | 8 (17) | 0.38

## Slide 13: Clinical Case Characteristics (n = 10)
HCU Exp. > 5 hours | CPB Time > 2 hours | Impaired Immune Fx. | CV Implant | Pleural Fluid | Blood | Spleen | Liver | Bone Marrow | Deep Sternal wound | Deep Abscess | NTM Organism
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
X | X | X | X |  | X | X |  | X |  |  | MAC NOS
X | X |  | X |  | X |  |  | X |  |  | M. chimaera
X | X |  | X |  |  |  |  | X |  |  | M. chimaera
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
. | X |  | X |  | X |  | X | X |  |  | M. chimaera
X | . |  | X |  |  |  |  |  |  | X | MAC NOS
. | . |  |  |  |  |  |  |  | X |  | MAC NOS
. | NA |  |  |  |  |  |  |  | X |  | NTM NOS
. | NA | X |  |  | X |  |  |  |  |  | Rapid growing NTM

## Slide 14: Clinical Case Characteristics (n = 10)
HCU Exp. > 5 hours | CPB Time > 2 hours | Impaired Immune Fx. | CV Implant | Pleural Fluid | Blood | Spleen | Liver | Bone Marrow | Deep Sternal wound | Deep Abscess | NTM Organism
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
X | X | X | X |  | X | X |  | X |  |  | MAC NOS
X | X |  | X |  | X |  |  | X |  |  | M. chimaera
X | X |  | X |  |  |  |  | X |  |  | M. chimaera
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
. | X |  | X |  | X |  | X | X |  |  | M. chimaera
X | . |  | X |  |  |  |  |  |  | X | MAC NOS
. | . |  |  |  |  |  |  |  | X |  | MAC NOS
. | NA |  |  |  |  |  |  |  | X |  | NTM NOS
. | NA | X |  |  | X |  |  |  |  |  | Rapid growing NTM

## Slide 15: Clinical Case Characteristics (n = 10)
HCU Exp. > 5 hours | CPB Time > 2 hours | Impaired Immune Fx. | CV Implant | Pleural Fluid | Blood | Spleen | Liver | Bone Marrow | Deep Sternal wound | Deep Abscess | NTM Organism
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
X | X | X | X |  | X | X |  | X |  |  | MAC NOS
X | X |  | X |  | X |  |  | X |  |  | M. chimaera
X | X |  | X |  |  |  |  | X |  |  | M. chimaera
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
. | X |  | X |  | X |  | X | X |  |  | M. chimaera
X | . |  | X |  |  |  |  |  |  | X | MAC NOS
. | . |  |  |  |  |  |  |  | X |  | MAC NOS
. | NA |  |  |  |  |  |  |  | X |  | NTM NOS
. | NA | X |  |  | X |  |  |  |  |  | Rapid growing NTM

## Slide 16: Clinical Case Characteristics (n = 10)
HCU Exp. > 5 hours | CPB Time > 2 hours | Impaired Immune Fx. | CV Implant | Pleural Fluid | Blood | Spleen | Liver | Bone Marrow | Deep Sternal wound | Deep Abscess | NTM Organism
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
X | X | X | X |  | X | X |  | X |  |  | MAC NOS
X | X |  | X |  | X |  |  | X |  |  | M. chimaera
X | X |  | X |  |  |  |  | X |  |  | M. chimaera
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
. | X |  | X |  | X |  | X | X |  |  | M. chimaera
X | . |  | X |  |  |  |  |  |  | X | MAC NOS
. | . |  |  |  |  |  |  |  | X |  | MAC NOS
. | NA |  |  |  |  |  |  |  | X |  | NTM NOS
. | NA | X |  |  | X |  |  |  |  |  | Rapid growing NTM

## Slide 17: Clinical Case Characteristics (n = 10)
HCU Exp. > 5 hours | CPB Time > 2 hours | Impaired Immune Fx. | CV Implant | Pleural Fluid | Blood | Spleen | Liver | Bone Marrow | Deep Sternal wound | Deep Abscess | NTM Organism
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
X | X | X | X |  | X | X |  | X |  |  | MAC NOS
X | X |  | X |  | X |  |  | X |  |  | M. chimaera
X | X |  | X |  |  |  |  | X |  |  | M. chimaera
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
. | X |  | X |  | X |  | X | X |  |  | M. chimaera
X | . |  | X |  |  |  |  |  |  | X | MAC NOS
. | . |  |  |  |  |  |  |  | X |  | MAC NOS
. | NA |  |  |  |  |  |  |  | X |  | NTM NOS
. | NA | X |  |  | X |  |  |  |  |  | Rapid growing NTM

## Slide 18: Clinical Case Characteristics (n = 10)
HCU Exp. > 5 hours | CPB Time > 2 hours | Impaired Immune Fx. | CV Implant | Pleural Fluid | Blood | Spleen | Liver | Bone Marrow | Deep Sternal wound | Deep Abscess | NTM Organism
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
X | X | X | X |  | X | X |  | X |  |  | MAC NOS
X | X |  | X |  | X |  |  | X |  |  | M. chimaera
X | X |  | X |  |  |  |  | X |  |  | M. chimaera
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
. | X |  | X |  | X |  | X | X |  |  | M. chimaera
X | . |  | X |  |  |  |  |  |  | X | MAC NOS
. | . |  |  |  |  |  |  |  | X |  | MAC NOS
. | NA |  |  |  |  |  |  |  | X |  | NTM NOS
. | NA | X |  |  | X |  |  |  |  |  | Rapid growing NTM

## Slide 19: Clinical Case Characteristics (n = 10)
HCU Exp. > 5 hours | CPB Time > 2 hours | Impaired Immune Fx. | CV Implant | Pleural Fluid | Blood | Spleen | Liver | Bone Marrow | Deep Sternal wound | Deep Abscess | NTM Organism
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
X | X | X | X |  | X | X |  | X |  |  | MAC NOS
X | X |  | X |  | X |  |  | X |  |  | M. chimaera
X | X |  | X |  |  |  |  | X |  |  | M. chimaera
X | X | X | X | X |  |  |  |  |  |  | MAC NOS
. | X |  | X |  | X |  | X | X |  |  | M. chimaera
X | . |  | X |  |  |  |  |  |  | X | MAC NOS
. | . |  |  |  |  |  |  |  | X |  | MAC NOS
. | NA |  |  |  |  |  |  |  | X |  | NTM NOS
. | NA | X |  |  | X |  |  |  |  |  | Rapid growing NTM

## Slide 20: Time between surgery and NTM+ Culture
Chart: Mean number of months
| 0-6 months | 7-12months | 13-18 months | 19-24 months | 25-30 months
 | 3 | 2 | 1 | 2 | 1

## Slide 21: Surgery Type and Odds of NTM Infection
Surgery Type | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Crude Odds Ratio (95% CI)
Any major CT surgery | 10 (100) | 31 (65) | 11.7 (0.6 – 211.3)
Major cardiac surgery1 | 8 (80) | 21 (44) | 5.1 (1.0 – 26.8)
Major thoracic surgery2 | 2 (20) | 11 (23) | 0.84 (0.2 – 4.6)
1Major cardiac surgery: CABG, valve replacement, aortic graft, pericardial window
2Major thoracic surgery: lung removal, thoracotomy, esophagectomy

## Slide 22: Surgery Characteristics and Odds of NTM Infection
Surgical exposures | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Odds Ratio (95% CI)
CBP | 8 (80) | 20 (42) | 5.6 (1.1 – 29.2)
Artificial valve or graft | 7 (70) | 9 (19) | 10.1 (2.2 – 46.9)
Topical medication | 9 (90) | 37 (77) | 2.7 (0.3 – 23.5)
Topical antibiotic | 9 (90) | 31 (65) | 4.9 (0.6 – 42.3)
Topical anticoagulant | 5 (50) | 15 (31) | 2.2 (0.6 – 8.8)
CBP = cardiopulmonary bypass

## Slide 23: Exposure to HCU and On-Pump Time
1Two-sample t-test for pooled variance
 | Probable Cases N = 10 minutes | Control Patients N = 48 minutes | P-value1
Cumulative HCU exposure |  |  | 
mean | 261.4 | 108.9 | 0.003
(IQR) | (164 – 333) | (0 – 221.5) | 
Cumulative time on CPB |  |  | 
mean | 117.4 | 37.8 | 0.0002
(IQR) | (46 – 169) | (0 – 73.5) | 

## Slide 24: Duration of Exposure and Odds of NTM Infection
 | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Odds Ratio (95% CI)
HCU exposure1 > 5 hours | 6 (60) | 5 (10) | 12.9 (2.7 – 61.9)
CPB on-pump > 2 hours | 6 (60) | 4 (8) | 16.5 (3.2 – 84.0)
Length of HCU exposure1 |  |  | 
0 hours | 2 (20) | 28 (58) | ref
> 0 – 4 hours | 1 (10) | 9 (19) | 1.6 (0.1 – 19.2)
> 4 – 5 hours | 1 (10) | 6 (13) | 2.3 (0.2 – 30.1)
> 5 hours | 6 (60) | 5 (10) | 16.8 (2.6 – 108.1)
1Cumulative operating time (OR entry and exit) with HCU present

## Slide 25: Duration of Exposure and Odds of NTM Infection
 | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Odds Ratio (95% CI)
HCU exposure1 > 5 hours | 6 (60) | 5 (10) | 12.9 (2.7 – 61.9)
CPB on-pump > 2 hours | 6 (60) | 4 (8) | 16.5 (3.2 – 84.0)
Length of HCU exposure1 |  |  | 
0 hours | 2 (20) | 28 (58) | ref
> 0 – 4 hours | 1 (10) | 9 (19) | 1.6 (0.1 – 19.2)
> 4 – 5 hours | 1 (10) | 6 (13) | 2.3 (0.2 – 30.1)
> 5 hours | 6 (60) | 5 (10) | 16.8 (2.6 – 108.1)
1Cumulative operating time (OR entry and exit) with HCU present
Non-MAC NTM

## Slide 26: Duration of Exposure and Odds of NTM Infection
 | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Odds Ratio (95% CI)
HCU exposure1 > 5 hours | 6 (60) | 5 (10) | 12.9 (2.7 – 61.9)
CPB on-pump > 2 hours | 6 (60) | 4 (8) | 16.5 (3.2 – 84.0)
Length of HCU exposure1 |  |  | 
0 hours | 2 (20) | 28 (58) | ref
> 0 – 4 hours | 1 (10) | 9 (19) | 1.6 (0.1 – 19.2)
> 4 – 5 hours | 1 (10) | 6 (13) | 2.3 (0.2 – 30.1)
> 5 hours | 6 (60) | 5 (10) | 16.8 (2.6 – 108.1)
1Cumulative operating time (OR entry and exit) with HCU present
Deep Sternal SSI

## Slide 27: Duration of Exposure and Odds of NTM Infection
 | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Odds Ratio (95% CI)
HCU exposure1 > 5 hours | 6 (60) | 5 (10) | 12.9 (2.7 – 61.9)
CPB on-pump > 2 hours | 6 (60) | 4 (8) | 16.5 (3.2 – 84.0)
Length of HCU exposure1 |  |  | 
0 hours | 2 (20) | 28 (58) | ref
> 0 – 4 hours | 1 (10) | 9 (19) | 1.6 (0.1 – 19.2)
> 4 – 5 hours | 1 (10) | 6 (13) | 2.3 (0.2 – 30.1)
> 5 hours | 6 (60) | 5 (10) | 16.8 (2.6 – 108.1)
1Cumulative operating time (OR entry and exit) with HCU present
Vertebral Osteo

## Slide 28: Duration of Exposure and Odds of NTM Infection
 | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Odds Ratio (95% CI)
HCU exposure1 > 5 hours | 6 (60) | 5 (10) | 12.9 (2.7 – 61.9)
CPB on-pump > 2 hours | 6 (60) | 4 (8) | 16.5 (3.2 – 84.0)
Length of HCU exposure1 |  |  | 
0 hours | 2 (20) | 28 (58) | ref
> 0 – 4 hours | 1 (10) | 9 (19) | 1.6 (0.1 – 19.2)
> 4 – 5 hours | 1 (10) | 6 (13) | 2.3 (0.2 – 30.1)
> 5 hours | 6 (60) | 5 (10) | 16.8 (2.6 – 108.1)
1Cumulative operating time (OR entry and exit) with HCU present
Pleural and Disseminated Dz

## Slide 29: Post-Operative Exposures
 | Probable Cases N = 10 n (%) | Control Pts. N = 48 n (%) | Odds Ratio (95% CI)
Central line | 9 (90) | 27 (56) | 7.0 (0.8 – 59.7)
Chest tube | 9 (90) | 31 (65) | 4.9 (0.6 – 42.3)
Shower before discharge | 8 (90) | 20 (42) | 5.6 (1.1 – 29.2)

## Slide 30: Operating Room Schematic
Air sample
Surface sample
Air sample
Water sample
Water sample
Water sample
Scrub sink
door
patient
CPB Device
HCU
Ice machine
CPB Device
OR Exhaust

## Slide 31: Water Samples
Air sample
Surface sample
Air sample
Scrub sink
door
patient
HCU
Water sample
Water sample
Ice machine
Water sample
CPB Device
OR Exhaust

## Slide 32: HCU Reservoir Swab
Air sample
Air sample
Scrub sink
door
patient
HCU
Water sample
Water sample
Surface sample
Ice machine
Water sample
CPB Device
OR Exhaust

## Slide 33: Large Volume Air Sampling
Scrub sink
door
patient
HCU
Air sample
Water sample
Water sample
Surface sample
Ice machine
Water sample
CPB Device
Air sample
OR Exhaust

## Slide 34: Available Isolates
56 environmental samples
  Water samples
  Surface samples
  Air samples
3 patient isolates
Scrub sink
door
patient
HCU
Air sample
Water sample
Water sample
Surface sample
Ice machine
Water sample
CPB Device
Air sample
OR Exhaust

## Slide 35: Available Isolates
56 environmental samples
  Water samples
  Surface samples
  Air samples
3 patient isolates
Scrub sink
door
patient
HCU
Air sample
Water sample
Water sample
Surface sample
Ice machine
Water sample
CPB Device
Air sample
OR Exhaust

## Slide 36: Simulated Procedure Air Sampling Results
Time | Bypass Status | 18” from HCU exhaust | Next to room exhaust
T0 | Standby | negative | negative
T5m | Standby | negative | negative
T1h | Cardioplegia on & circulatory arrest | negative | negative
T2h | Cardioplegia on | M. chimaera | AFB +, not NTM
T3h | Warming | M. chimaera | negative
T4h | Warming | M. chimaera | negative
T5h + 10m | Off for 10 minutes | negative | AFB +, not NTM

## Slide 37: HCU Internal Water Pathway

## Slide 38: Water Pathway Laboratory Results
Sample | Result
HCU 1 (shipped to CDC) | 
Water (drained by facility) | 15 CFU AFB +
Water (drained after shipping; machine running) | > 3.0 x 104 CFU M. chimaera
Swab of biofilm inside reservoir | M. chimaera
HCU 2 (simulation) | 
Water (drained by facility) | 10 CFU AFB +
Water (drained immediately after simulation) | 1.5 x 105 M. chimaera
Swab | Not performed
HCU 3 | 
Water (drained by facility) | 300 CFU AFB +
Water (drained from static machine) | 140 CFU AFB +
Swab of biofilm inside reservoir | M. chimaera
Cardioplegia machine | 
Water circuit interior | M. chimaera

## Slide 39: Water Pathway Laboratory Results
Sample | Result
HCU 1 (shipped to CDC) | 
Water (drained by facility) | 15 CFU AFB +
Water (drained after shipping; machine running) | > 3.0 x 104 CFU M. chimaera
Swab of biofilm inside reservoir | M. chimaera
HCU 2 (simulation) | 
Water (drained by facility) | 10 CFU AFB +
Water (drained immediately after simulation) | 1.5 x 105 M. chimaera
Swab | Not performed
HCU 3 | 
Water (drained by facility) | 300 CFU AFB +
Water (drained from static machine) | 140 CFU AFB +
Swab of biofilm inside reservoir | M. chimaera
Cardioplegia machine | 
Water circuit interior | M. chimaera

## Slide 40: Water Pathway Laboratory Results
Sample | Result
HCU 1 (shipped to CDC) | 
Water (drained by facility) | 15 CFU AFB +
Water (drained after shipping; machine running) | > 3.0 x 104 CFU M. chimaera
Swab of biofilm inside reservoir | M. chimaera
HCU 2 (simulation) | 
Water (drained by facility) | 10 CFU AFB +
Water (drained immediately after simulation) | 1.5 x 105 M. chimaera
Swab | Not performed
HCU 3 | 
Water (drained by facility) | 300 CFU AFB +
Water (drained from static machine) | 140 CFU AFB +
Swab of biofilm inside reservoir | M. chimaera
Cardioplegia machine | 
Water circuit interior | M. chimaera

## Slide 41: Isolate Fingerprinting — Highly Related PFGE Patterns
  Clinical cultures of M. chimaera from 3 case patients
  Environmental M. chimaera isolates
Possible pseudo-infection: isolate from native valve explanted due to non-NTM endocarditis

## Slide 42: Limitations
Experience of one facility
  Snapshot in time
  Small sample size, only 10 case-patients
Epi and genetic diversity of M. chimaera not well established

## Slide 43: Investigation Summary
Confirmed outbreak of extrapulmonary MAC infections of several years duration
Found epidemiological association between infection, cardiac surgery, CPB and HCU exposure; findings supported by laboratory investigation
Observed long periods of time between procedure and symptom onset (months to years)
Identified significant challenges to precisely follow IFU; multiple versions of IFU noted
Observed HCU contamination with biofilm and demonstrated that HCU can disperse M. chimaera

## Slide 44: Increasing Awareness: Prevention and Case Finding
FDA Safety Communication
CDC Interim Practical Guidance
Initial PADOH Health Alert Network (HAN) Alert
“Look and you will find it —
what is unsought will go undetected.”
— Sophocles

## Slide 45: Turbid Water

## Slide 46: Secondary Overflow Tube

## Slide 47: Drain Tube and Plug

## Slide 48: Many Hard to Clean Connections

## Slide 49: Filling HCUs and Point-of-Use Filters

## Slide 50: Active Cooling

## Slide 51: Ice Bath

## Slide 52: How is ice made? Where is it used? Is the water filtered?

## Slide 53: Detailed Guidance to Facilities Issued through PA HAN
http://www.health.pa.gov/Your-Department-of-Health/Offices%20and%20Bureaus/epidemiology/Pages/2015-Health-Alerts.aspx#.VyweJhUrKAw

## Slide 54: Minimize Risk but Don’t Delay Surgery
Patients and providers should not delay surgery after appropriate informed consent
In the absence of additional information, follow IFU
Ensure facility is using most current version of IFU
This is a patient safety issue, use a comprehensive approach to risk assessment and mitigation

## Slide 55: Comprehensive Approach
Comprehensive approach to reduce risk of bacterial colonization and subsequent aerosolization
  Assess physical device, ergonomics, clinical environment, individual, team and organizational behaviors
  Implement interventions using a hierarchy of controls (elimination/substitution, engineering controls, and administrative controls)

## Slide 56: Assessment
Identify devices and stakeholders
Assess status of devices, if possibly visibly inspect
Evaluate interplay of connected devices (reservoir?)
Assess SOPs relative to mfg’s IFU
Identify expertise needed for proper implementation
Assess training and existing competencies
Evaluate existing workflow to identify challenges and available resources
Evaluate impact of any changes

## Slide 57: Administrative Controls
Implement maintenance, disinfect and replacement schedule according to strictest interpretation of IFU
Implement process to ensure use of current IFU
Verify appropriateness and concentration of additives
Implement training program and develop job action sheets
Implement QA program to reassess interventions

## Slide 58: Elimination, Process and Engineering Controls
Do not use tap water, use a 0.2 micron filter. Only use sterile water, distilled water or deionized water if consistent with mfg’s IFU
Implement safeguards to minimize spillage and inappropriate “topping-off”
Move HCU away from surgical field (be thoughtful)

## Slide 59: Bacteriologic Surveillance
  Log device and on-pump time for each patient
  Implement a water quality sampling program
    Follow IFU, if not specified by manufacturer…
      Conduct monthly heterotrophic plate counts (HPC)
      At a minimum, water should meet EPA drinking water standards
      Consult with PADOH on NTM specific cultures and air sampling if a cluster of infections is suspected or if device assessments suggest contamination

## Slide 60: Patient Surveillance
Disseminate information to raise awareness among clinical staff
Ensure access to equipment and procedures to culture NTM in a patient with suspected disease
Conduct prospective and retrospective surveillance using proposed case definitions
Cross-reference microbiology NTM culture data with operative databases (important to note that many infections will not meet National Health Care Safety Network criteria)

## Slide 61: PADOH Interim Surveillance Case Definition Summary
 |  | Criteria
Timeframe and Exposure History |  | Onset of symptoms occur 2 weeks to 4 years after having had an open heart surgery on cardiopulmonary bypass
Infection |  | 
OR | Localized | Prosthetic valve Graft Sternotomy wound
 | Disseminated | Bacteremia Embolic and immunologic manifestations
Specimen |  | Bacteria identified in an invasive sample (NOT via bronchoscopy)
Pathology Result |  | Mycobacterium chimaera Mycobacterium avium complex (MAC) Mycobacterium-avium-intracellulare-scrofulaceum (MAIS)
 |  | Non-tuberculous Mycobacterium OR Consistent histopathology
Probable Case
Suspect Case

## Slide 62: Reporting and Taking Machines Out-of-Service
  At this time, PADOH cannot categorically recommend the elimination or substitution of one HCU over another unless there is evidence (visible, bacteriologic or epidemiologic) to suggest bacterial colonization or transmission
However, HCUs with discoloration, visible biofilm, visibly turbid water, or machines with water cultures outside of EPA limits for drinking water should be sequestered until further directed by PADOH
Notify FDA through MedWatch and manufacturer
Report these findings, departures from IFU and identified cases to PADOH

## Slide 63: Impact — Additional HCU Investigations
Four hospitals reported possible HCU contamination but no cluster of probable cases
Aware of five hospitals that have modified their existing HCUs or made or intend to make HCU purchasing decisions
Anecdotal reports that it is very difficult to achieve HPC levels that meet EPA drinking water standards in HCUs from a variety of manufacturers

## Slide 64: Impact — Patients
Two hospitals reported possible HCU contamination and initially identified 11 probable cases
~ 3,700 patients notified of possible exposure

## Slide 65: Notification of Patients in PA
Best practices provided by PADOH and CDC Patient Notification Toolkit; however, patient notification executed by hospital
Direct communication to identified case patients
Notification letters to potentially exposed patients and their primary care providers
Call center: must provide empathy and address clinical and financial (e.g., billing) aspects of care
Transition to dedicated NTM clinic

## Slide 66: Patient Notification: Clinical Challenges
Lack of sensitive laboratory screening tests
No prophylactic antibiotics
Vague presentation of illness
Provider may need to order AFB-specific culture, sensitivity likely improved with specimen collection in isolator culture tubes
Laboratory must use special techniques to isolate slow growing NTM, M. chimaera identification requires additional work

## Slide 67: Improving Guidance for Facilities — HCU
Does my HCU have a problem (retrospective and prospective surveillance)?
But… I have an HCU from manufacturer X?
Are current IFU adequate (and adequately validated)? Can I do more? Any disadvantages?
What should I do with existing machines?
I think I fixed my problem. How do we put machines back in service?
What do I do with attached devices?

## Slide 68: Improving Guidance for Public Health and Providers
Appropriate case definition?
  Up to 4 years after exposure or more?
  MAC and M. chimaera only or other NTM?
Does early diagnosis change patient outcome?
Optimal strategy of screening and monitoring asymptomatic exposed patients?
Optimal medical and surgical management once diagnosed?

## Slide 69: Conclusions
Multiple studies have now demonstrated an epidemiological association between NTM infection and HCUs. These findings are supported by laboratory investigations.
Risk of NTM transmission appears real, with severe consequences, though infections appear to occur infrequently.
HCUs enable life-saving cardiac surgery; it should not be delayed.
Communication with patients and providers to promote informed decision making, appropriate vigilance and early detection is critical.
Clinical options are sub-optimal. More research is needed to guide patient management.
Greater transparency, and possibly more research, is needed to adequately assess and mitigate risk related to the propensity of HCUs to become contaminated, colonized and generate bacterial aerosols. Similarly, more information is needed to assess the adequacy of disinfection procedures to prevent transmission.

## Slide 70: Thank You!
Team at the Pennsylvania Department of Health
Colleagues at FDA and CDC
Hospitals and dedicated health care providers in PA
FDA for opportunity to discuss this important topic

## Slide 71: Questions
c-jeffmil@pa.gov
The findings and conclusions in this report are those of the authors and do not necessarily represent the official position of the Pennsylvania Department of Health, Centers for Disease Control and Prevention or the United States Public Health Service.